Cassini Basic Training Seminar
Developing and Running Test Plans

---
Seminar Outline
Cassini Basic Seminar Outline
- Operation and Troubleshooting
- System Administration and Maintenance
- Developing and Running RI Test Plans
- Science of RF Measurement
- Device Definitions
- Example Applications Development
- Test Fixture and Device Interface Design
- Test Design & Best Practices Test Optimization
- Application User Guides

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Test Plan Concepts
States
Measurements
Limits
Inheritance - global defaults, section defaults
Buttons (objects)
Run time vs. Design time
[Speaker Notes: grafical user interface where a designer can assemble tests by working with buttons instead of writing source code. The grafical test plans are compiled into optimized machine code that is used to control the instruments in the system.]

---
Test Plan Structure

---
Typical Test Plan
[Speaker Notes: grafical user interface where a designer can assemble tests by working with buttons instead of writing source code. The grafical test plans are compiled into optimized machine code that is used to control the instruments in the system.]

---
Creating Test Plan
Creating a New Test Plan
Change the Device Type
Change the Interface Board
Change the Fixture
Edit the new Test Plan

---
Creating a New Test Plan
From the RI Message Window, select Test and Plans...
From the Select Testplan window, select New

---
Save and Title the Test Plan
File > Save Guru

---
Selecting the Device Type
From Tester menu, Change dut
Select the device
Repeat for all Tester
resources (DUT, DUT control, DIB, and Fixture)

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Test Plan Structure
Global Defaults
Disconnect Settings
Connect Sequence
Adding a Test Section
Conditional Statement
Section Defaults
Adding Test to a Test Section

---
New Test Plan
Name of the test plan is displayed in the title bar
Test plan outline panel and setting panel
[Speaker Notes: Talk about test plan outline panel and test plan setting panel The first line in the test plan outline is the test plan settings line or test plan header. The test plan header does not have a test/settings panel associated with it. If the Test Plan Settings line is selected [highlighted] in the test plan outline, the working space shown in the test/settings panel will always be blank. The first three lines of the new test plan outline is displayed in the test plan outline panel as default]

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Global Defaults Panel
Global Defaults define the Test Plan's default settings for all of the instruments in the tester.

---
Disconnect Settings Panel
Defines the specific state of the instruments at the moment the device-under-test is disconnected from the tester

---
Connect Sequence Panel
Defines the order in which settings are activated
Order of execution is from left to right

---
Adding a Test Section
Each test plan must have at least one test section
Select menu Edit & Add Test Section

---
Test Section Title
Enter the title of the new test section.
[Speaker Notes: The System will display the dialog box: Title for New Test Section? OK button or press the Enter when finish]

---
Test Section Layout
Test Section Header
Conditional Statement
Section Defaults
One or more Test
[Speaker Notes: Test sections enable you to group together individual tests conditional statements; logical statements for deciding if the tests in this test section should be performed or skipped Section Defaults define the default instrument settings for the tests in this test section Test Plan Optimizer, only optimize within a test section]

---
Adding Test to a Test Section
Select Test Section Header line in the test plan outline panel.
Select menu Edit and Add Test.

---
Test Title
Enter the title for the new test.
[Speaker Notes: System displays the Title for New Test? dialog box. Click OK or press Enter when finish]

---
Additional Tests and Test Sections
Using the same procedure we have just describe

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Building a Test
Open the Tester Viewer window
"Select" to Copy the Desired Button
Paste the Button into the Test Panel
Organizing the Test Panel

---
Open the Tester Viewer window
Select menu Tester and View

---
"Select" to Copy the Desired Button
Select a instrument
Select measurement or state
RMBC (mouse button 2) on the button then click Select from the pop-up menu

---
Selected Button is Highlighted
Tester Viewer window will indicate selected button by highlighting the button.

---
Paste the Button into the Test Panel
Change the active window to the Test Plan Editor
RMBC (mouse button 2) on the desired location and select Paste Button from the pop-up menu.

---
Organizing the Test Panel
RMBC (mouse button 2) and hold on the desired button then drag the button to new location
Copying, Pasting, and Moving Buttons in the same process
Short cut on copy button
Ctrl + LMBC (mouse button 1)
Short cut on paste button
Ctrl + RMBC (mouse button 2)

---
Changing from Single Value to Range
Place the mouse pointer inside a state button, select pop-up menu Range

---
Grouped Buttons

---
Modifying Button Settings

---
Modifying Button Settings
Type in the new value then click OK or press Enter

---
Test Example: RF Gain
[Speaker Notes: In the example shown, the vna will perform the GAIN measurement for 24 different conditions (6 frequency x 4 voltage levels x 1 power level)]

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Saving Device Data
Add the SAVE DATA buttons to the test panel
Connect the MEAS button to the SAVE DATA button
Each SAVE DATA button in the test plan must have a unique data name

---
Adding the SAVE DATA button
Go to the Tester Viewer window.
Select the System choice from the Instrument list.
Select the data saves choice from the Button Type list.
Select and copy one of the SAVE DATA buttons:
SAVE NO FORMAT
SAVE FORMAT
SAVE ADJUSTED
Return to the Test Plan Editor and paste
the Return to the Test Plan Editor and paste
the SAVE DATA button button

---
Data Name for the SAVE DATA button
RMBC on the SAVE DATA button
Select Data Name
Enter the desired Data Name

---
Connecting MEAS & SAVE DATA buttons
Press and hold RMB the sending data port of the MEAS button,
Drag to the data receive port of SAVE DATA
To disconnect the button
repeat the process again.

---
Adding CALC button
Insert a general purpose CALC button and connect the lines as shown above.

---
Setting the CALC button
RMBC and select Calculation...
Select a calculation

---
Saving Calibration Data
Connect the CALC button to the another CALC button, LOCAL VAR SOURCE, or
SAVE FORMAT buttons
A sending data port
can connect to
zero or more
data receive port

---
LOCAL VAR buttons
LOCAL VAR SAVE buttons
LOCAL VAR SOURCE buttons
Use a LOCAL VAR SAVE button with a
LOCAL VAR SOURCE button
to transfer test results from
one test panel to another

---
Selecting the Data Format
Specify the data format (units) used by the system to save the data and limit check the data
RMBC on SAVE Data button,
select Data Form from the pop-up menu
select one of format choice from the pop-up menu

---
Data Format: Index, Linear, Log

---
Tester Configuration for the Test Plan
Physical Tester
Loaded Tester
View Tester Resources
[Speaker Notes: Define each term for the student. Physical, Target, Edit, Active tester Typically, you create a test plan with the Edit Tester and you compile your test plan against the Active Tester In most cast Edit Tester is the same as the Target Tester]

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Tester System Mode
ACTIVE HARWARE MODE!!!!
SIMULATION MODE!!!!

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Compiling & Running Test Plan
Before a test plan can be executed, we need to compile and optimize it for the Tester
Be sure that the current Active tester is the physical tester configuration

---
Compile and Optimize Test Plan
From the Test Plan Editor, select Test Plan menu & Compile
Or the press Compile button on the top right

---
Compile Errors
If an error is discovered during the compile process
Details are displayed in the RI Message Window

---
Running a Test Plan
Select Test Plan menu & Run
or just press Run button on the top right

---
Time Runs for a Test Plan
To time how fast a Test Plan runs
Select Options menu & Time Runs
[Speaker Notes: Time Runs is a on/off button. Once turn on, on every press of run, the test time will display on the System Message Window This is a good esterment, Text Exec will give the most acure time]

---
System Message window - Timing
Detail timing display on the System Message window

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Viewing Test Results
Data Viewer displays the SAVE DATA button's Name and the name of the instrument settings button in its title bar
Set the parameters, formats and order
[Speaker Notes: SAVE DATA button's Name and the name of the instrument settings button in its title bar test results are displayed in the test results plot in the middle of the viewer. select and set the parameters, formats and order in which the data is to be presented with the Data Viewer buttons provided. Data Viewer will set its SORT BY parameter and the X AXIS parameter equal to the name of the instrument settings button Data Viewer will also set its X axis minimum and maximum values equal to the start and stop points in the instrument settings button.]

---
Viewing Test Data
RMBC on SAVE DATA button, select View Data from the pop-up menu

---
Select Viewer Type
Select a list of data viewer choices from the
Select Viewer dialog box
Rectangular
Smith Chart
Polar
Strip Chart

---
On Screen Help Text
RMBC on any of the button, select Show Note...
For extensive on-screen reference information, click on Help menu

---
Change Data Layout Format (SORT BY)
PARAMETER button displays the parameter that the data is sorted by
FORMAT button is used to displays the order of format in which the data is presented
[Speaker Notes: SORT BY PARAMETER and SORT BY FORMAT buttons are on the left side of the data viewer panel and are used to select and display the sorting criteria for the data.]

---
Change Data Viewer Sort Parameters
Y PARAMETER button select the measurement parameter data you want displayed along the Y axis

---
Change Data Viewer Format
Y FORMAT button select the units and format of the Y Axis data.

---
Format Parameters
Formats available depend on the measurement parameter selected
Ex. for the S21 measurement

---
Change Data Viewer Axis Scale
MIN X and MAX X
FORMAT X axis
Y CENTER and Y SCALE

---
Change Data Viewer Normalization
NORMALIZE button enable you to save the test results currently displayed on the test results plot and to display succeeding test result traces relative to the test results data saved

---
OverPaint Data Viewer Option
OverPaint is a toggle function

---
Using OverPaint Viewer Option
RMBC on the Data Viewer and select OverPaint from the pop-up menu

---
Adding Memo to Data Viewer
MEMO button provides capability to add text or comments to each plot
Click on the MENO button and enter the text

---
Use Data Viewer to Identify Data Point
LMBC on the Data Viewer identify the X and Y coordinates

---
Delta Coordinates
Identify the difference between two point by click and hold on LMB on the first point then drag to the second point
(delta X & delta Y coordinates)

---
Data Viewer Mouse Button 2 Functions
RMBC on the Data Viewer
Clear Traces
Redraw
Trace 1 color
Grid color
Lines
Marker
Clear
Max
Min
X select
Line Type

---
Multi-Dimentional Data Viewers
Multi-Dimensional Data Viwer uses different colored traces to distinguish between different parameter values

---
Opening a Multi-Dimentional Data Viewer
RMBC on SAVE DATA button, select the
View Data option
Select parameter for the X axis
Select parameter to
display different
colored trace lines
[Speaker Notes: Set each of these remaining parameters equal to a constatn by selecting [highlighting[ a test point Value associated with each parameter. You can ony have two intependent variables displayed at one time. The remaining parameters must be set to a constant value.]

---
Select a Viewer
Select a viewer from the Select Viewer dialog box

---
Using the Multi-Dimensional Data Viewer
Multi-Dimensional Data Viewer similar to the
X & Y coordinate data viewer with added buttons color bar and Redraw
[Speaker Notes: color bar button is used to select and display the results vs x-axis To display all of the traces on the screen simultaneously, use the OverPaint (On) option and select each of the different colors in the color bar button or select the Redraw button which will redraw all of the test result traces in sequential order.]

---
Setting and Selecting Limits
Each test plan can have one or more sets of test limits associated with the test plan.
Need compiled and run the test plan at least once before creating test limits
To set test limits to the test plan, you must first either open an existing set of test limits or create a new set of test limits.

---
Create new set of test limits
Select the Test Plan Editors menu Limits and New

---
Histogram
Enter a title for the new test limits and click OK
Reserved name Histogram
[Speaker Notes: Reserved name Histogram, Verifacation A, Vererifacation B, and AOQ]

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Selecting Test Limits
Select the Test Plan Editors menu Limits and Select...

---
Test Limits
Test Plan Editor displays the Select Test Limits dialog box
Select the desired test limits

---
Adding Single Valued Test Limits
RMBC on the DATA SAVE button
Select the Single-valued Limit from pop-up menu

---
Test Limit Types
Test Plan Editor will display the Set Test Limits dialog box
Enter the upper and/or lower limit
Select the OK button

---
Adding Multi-Valued Test Limits
RMBC on the DATA SAVE button
Select the Multi-Valued Limit from pop-up menu

---
Upper or Lower Limits
Enter the upper and/or lower limits for each of the point
[Speaker Notes: The test system only performs the test limit checks for those upper or lower test limits that have a value. Blank limit values are ignored.]

---
Saving the Test Limit
Save the test limits with the test plan
Test Plan Editors menu File and Save

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Release for Production
Change status to released
[Speaker Notes: After you have created, run, debugged, and created limits for the test plan, you can perform the release process. When you release a test plan, the System Controller flags the file as "released" and time stamps the file.]

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- ...
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production
- Opening a Previous Version (Guru)
- Package Exec: Handler Bin Defintions

---
Opening a Previous Version
Show Test Objects navigation window
[Speaker Notes: After you have created, run, debugged, and created limits for the test plan, you can perform the release process. When you release a test plan, the System Controller flags the file as "released" and time stamps the file.]

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- ...
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production
- Opening a Previous Version (Guru)
- Package Exec: Handler Bin Defintions

---
Package Exec: Handler Bins
Choose "hard" or "soft" bin for each category
Hard bin is the sort value sent to the handler
Soft bin is stored in data logs (STDF)
If using Retset Criteria, you should use a "re-plunge" capable hander
Set the "retset bin" signal to cause a handler re-plunge
[Speaker Notes: After you have created, run, debugged, and created limits for the test plan, you can perform the release process. When you release a test plan, the System Controller flags the file as "released" and time stamps the file.]

---
Package Exec: Handler Bins
[Speaker Notes: After you have created, run, debugged, and created limits for the test plan, you can perform the release process. When you release a test plan, the System Controller flags the file as "released" and time stamps the file.]

---
Chapter Outline (Ch 3)
- Creating a New Test Plan
- ...
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production
- Opening a Previous Version (Guru)
- Package Exec: Handler Bin Defintions

---
Chapter Review (Ch 3)
- Creating a New Test Plan
- Understanding Test Plan Structure
- Building a Test Plan
- Saving Data
- Editing the Tester Configuration
- Compiling & Running Test Plan
- Viewing Test Results
- Setting Limits
- Release Test Plan for Production

---
Next Topics Preview
- Science of RF Measurement
- Device Definitions
- Example Application Development
- Fixture ad Device Interface Design
- Test Design & Best Practices

---
Questions?
Any Questions from this Chapter?