CASSINI RF/ Microwave ATE System
Test Design &
Best Practice Test Optimization

---
Guidelines
Test Plan - Comments, Guidelines and Suggestions
Optimizer - Optimizes over Test Section
Only add state buttons that are needed
Changes cost time and money
Data is displayed in the order listed & not displayed in the order performed
Parameters set by LVs must be reset in the Disconnect Settings panel

---
Optimizer: Optimizes over each Test Section
Compiled Delta Settings Shows Test Times

---
Optimizer: Optimizes each Test Section.
Lowest DUT Frequency Measured First

---
Optimizer: Optimizes each Test Section.
Measurement with Lowest Input Power Performed First

---
Test Plan - Comments, Guidelines and Suggestions
The tester is very fast changing state
Your part may not be fast, add Pauses
The tester will find unstable DUT states

---
Test Plan - Comments, Guidelines and Suggestions
Fewer Test Sections better than more
More simple test panels are better than fewer complex test panels
Duplicate test states/conditions when ever possible
Be consistent about where you place buttons in a test panel

---
Test Plan - Comments, Guidelines and Suggestions
IMD Tests - more Receive Attn is better
Noise Figure - Use 0 dB Rec Attn.
Set DB Current Limit = DB I Meas Max.
Set VI Current Limit > VI I Meas Max.
RF Sources - Levels <-22 to >+13 dBm
DMSG - Optimum levels <-5 dBm
RF Off is <-30 dBm

---
System Receive Gain Adjustment
Receive Attenuation
IF Gain
Signal Dependent
 Single tone
 Multi tone
Measurement Repeatability

---
Simplified Receive Diagram
Receive Attenuator
Mixer
IF Filter
IF Amplifier
Receiver Detector

---
System Sensitivities
RI 7100A Receiver Mixer
 Saturate (Compression)
 Generate Intermodulation Distortion
Receiver Detector
 Compression
 "Fold-over"

---
Single Tone Measurements
RF Power, S-Parameters, Noise Figure
Objectives
 Maximize Power to Receiver
 Maximize Power to Mixer
 Set Appropriate IF Gain

---
Single Tone Method
Set IF Gain to 20 dB or Lower
Reduce Rx Atten Until Error Message
Backoff One Rx Atten Setting
Increase IF Gain Until Error Message
Backoff One IF Gain Setting

---
Multi-Tone Measurements
IM3/5/7 and ACPR
Measure Device not Tester's Mixer
Operate in the Mixer's Linear Range
The Mixer Can Add to or Subtract from the Measurement

---
Mixer Intermods
Receive Attenuator Controls Mixer IM Performance and System Noise Floor

---
Multi-Tone Method
Measuring Main Tone(s)
Not Sensitive to Mixer IMs
Use Method for Single Tone
Measuring IM & ACPR Tones
Sensitive to Mixer's IM performance
IF Gain usually at or near Maximum (Minimize noise floor)
Verify with Spectrum Analyzer (SA)
Be careful of SA's intermods!

---
Measure IM Tone
Set IF Gain 50/55; Receive Atten 20 dB
Measure C/I
Verify with Spectrum Analyzer
Go to break point
Set Receive input to Aux out
Measure C/I
If Different, Increase RF Attenuator
Repeat

---
IM Considerations
Verify High and Low IM Tones
Use Highest Linear Power DUT Condition
Typically 20 dB Rx Atten, 55 dB IF Gain
Too Much Rx Atten Raises Noise Floor

---
Effective Receiver Bandwidth

---
Organizing Your Test Plan
Define 1st RF state in Global Defaults
Add Fixture ON to Global Defaults
Define general test conditions in Section Default Panels
Define DC tests in 1st or last Section which require the RF Sources to be off
Define DC tests in 1st or last Section which require the DUT to be turned off

---
Organizing Your Test Plan
Define special startup sequences in the Connect Sequence panel or in the 1st Test Section
Write the tests in the order listed on your test list

---
Speeding-up Your Test Plan
VCC Changes & DC Measurements: < 2 msec.
RF Measurements: <30 msec.
RF Frequency Moves:	<20 msec. (Fast Settle) or	<1.5 msec (Freq Offset)
RF Mechanical Switches: >30 msec. & Expensive
Minimize RF Source and Receiver Frequency Moves
Minimize/Eliminate RF Mechanical Switch Changes:
 Changing RF Ports from Src to Rec or Rec to Src
 Changing RF Src 1 and RF Src 2 RF Attenuators
Using the Wide IF Filter is faster than Narrow IF Typically use 16 or less Averages for RF Measure
Typically use 32 or less Averages for DC Measure

---
Speeding-up Your Test Plan 2
Fast RF Meas: S parameters, RF Power and Phase
Slow RF Meas: ACPR, Noise Figure, IP3, Harmonics
Setting a Parameter to a Local Variable Value is very	costly : Requires a Mini-Compile
The Optimizer only Optimizes within a Test Section
Time Domain vs Freq Domain:
 If the RF Receiver can make the Measurement, it is
 faster with better Frequency & Dynamic Range
Oscope is designed to measure Pulse Characteristics
   Rise/Fall Time, Period, Duty Cycle, etc.
Receiver is designed to measure Complex Signals:
   Amplitude, Phase, Noise, Power, Frequency, etc.

---
Speeding-up Your Test Plan 3
Items to Set in Global Defaults
Receiver Freq Tracking to Source 1 or System
Test Head Parameter to b2
RF Ports to Receive or Source
Receiver IF Gain
Source 1 Freq & Level to first RF Stimulus Signal
Source 2 Freq & Level to the first IM Stimulus Signal
Source 3 Freq & Level to the first LO Stimulus Signal
Turning On and Off a RF Source is fast
Switching Source 1 Mode to Noise is faster!
Use Source 1 Mode button to Isolate Src 1 & 2 from DUT

---
Speeding-up Your Test Plan 4
Check Compiled Delta Settings to Improve Test Speed
Look for changes in Receiver & Source Frequency,	Attenuator & Test Head Settings, etc.
Minimize Test State changes by making test conditions	exactly the same for as many meas as possible:
i.e. RF Stimulus Levels, Frequencies, Receiver Settings, DC Levels, Current Limits, Ranges, etc.
Use b1 to test LO Leakage at the DUT's RF Input
Use Source 1 "back door" for S22 Measurements
For IM measurements, leave RF Source 1 at the normal stimulus Freq, this saves 2 RF Freq Moves
Remember DC only measurements still set the Receiver and RF Src's to their Default Frequencies

---
Typical RI 7100A RF ATE Measurement Time
Measurement	Time	Std. Dev
(msec.)
Power	2.5	<0.02 dB
Frequency	11.5	<0.5 Hz
S11, S12, S21 & S22	12	0.02 dB (S21)
Noise Figure	26	0.1 dB
Phase Noise	60	0.3 dB
IIP3	78	0.3 dB
P1dB	17.5	0.1 dB
ACLR or ACPR	93	0.4 dB
Harmonic Power	56	0.05 dB
I/Q Mod Phase & Amp	25	0.1 dB, 0.1 deg.
I/Q Demod Phase & Amp	60	0.02 dB, 0.15 deg.
DC Voltage	2	0.025%
DC Current	2	0.03%

---
Typical RI 7100A Noise Floor
<-155 dBm/Hz to 6 GHz

---
Typical Test System IIP3
+49 dBm
3 Testers (30 dB Rec Attenuation)

---
1 dB Noise Figure – Two Testers

---
Noise Floor Measurements
Pout +34.9 dBm, Noise Floor <-85 dBm

---
VCO Phase Noise Measurements
-114 dBc/Hz at 100 kHz Offset
-137 dBc/Hz at 1 MHz Offset

---
Production Measurement Performance Limits
DUT	Measurement	Performance
LNA	Noise Figure	<1.0 dB	IIP3	>+25 dBm
PA	CDMA ACPR1	<-60 dBc	CDMA ACPR2	<-70 dBc
Mixer	Noise Floor	<-150 dBm/Hz
PLL/VCO	Phase Noise	<-114 dBc/Hz @ 100 kHz
<-137 dBc/Hz @ 1 MHz

---
Production Measurement Performance Limits cont.
DUT	Measurement	Performance
Dividers	Frequency	<1 Hz Resolution to 20 GHz
I/Q Mod	Phase Match	<0.1 degree
Amplitude Match	<0.025%
LO Rejection	<-60 dBS
Image Rejection	<-70 dBS
I/Q Demod	Phase Match	<0.1 degree
Amplitude Match	<0.1 dB

---
Debugging and Breakpointing
Measurement Order and Test Flow
Setting a Breakpoint

---
Measurement Order and Test Flow
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Test Section Optimization
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Multiple States in a Panel
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Copy a Test or Test Section
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Setting a Breakpoint
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
RI Interactive Control Window
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Interactive DC Measurements
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Interactive RF Measurements
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Receiver Measurement Types
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Test Disable Usage
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Data Viewers
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Single Variable Data Viewers
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Overpaint Function
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Single Event Repetitive Measurements
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Moving Strip Chart Display
[Speaker Notes: Set the Hardware Bin to which different categories of tested parts will be placed by the handler Delay time, number of mSec system waits until it looks for a start of test from the handler Maximum wait time is number of sec system wait for start of test, before prompt the operator for actions]

---
Debugging Tips
Check active Calibration data for anomalies
Check each measurement for optimal RF Attenuator and IF Gain settings ( max S/N ).
Measurement must "Fit" in the Receiver's 70 dB instantaneous Dynamic Range
Small changes ( 1 to 5 dB ) in Device Power Level can be large if using all the dynamic range.	(i.e. measure a 50 to 60 dB carrier to tone.)
Setting resolution in the RF Attenuator is 10 dB steps and the IF Gain is 6 dB
[Speaker Notes: RI 7100 System software is a multi-user system running on a very stable operating system, OS/2 RI 7100 System software, AKA Test Executive. Regular administrative tasks as list here... Regular maintenance is recommended to assure continued operation, to increase MTBF]