Cassini RF ATE System
Basic Training Seminar
Science of RF Measurements

---
Seminar Outline
Cassini Basic Seminar Outline
- Operation and Troubleshooting
- System Administration and Maintenance
- Basic Test Plan Concepts
- Science of RF Measurement
- Device Definitions
- Example Applications Development
- Test Fixture and Device Interface Design
- Test Design & Best Practices Test Optimization
- Application User Guides

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Cassini ATE System
Basic Measurements
Intermod Distortion
3rd Order Intercept (TOI)
Isolation
DC Voltage & DC Current
DC to RF Efficiency
I & Q Amplitude & Phase
Rise/Fall, Duty Cycle, Period & Jitter
S Parameters
RF Power
Noise Figure &	Noise Power
Spectral Purity & ACP
Gain Compression
S/N, SINAD,	Distortion & BER

---
Verification
System startup verifies major component functionality
System receiver self calibration/verification every 20 minutes
Full verification to calibration specifications automatically run on user schedule or on demand

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Why Measure Phase and Magnitude ?
To Measure Complex Impedance
Resistance with Capacitance & Inductive components Complete Characterization of Linear Networks Transmission Distortion Effects
Deviation from Linear Phase
Bandwidth
Vector Accuracy Enhancement
De-embedding of measurement results
Characterization of Systematic Errors
Time Domain Transforms

---
Vector Network Measurement
Reflection Meas:
Input VSWR
Return Loss
Input Impedance
Reflection Coeff
S Parameters S11 S22
Transmission Meas:
Gain/Loss
Insertion Phase
Group Delay
S Parameters S21 S12

---
Typical Network Analyzer

---
RF Reflected Signals

---
Reflection Coefficient,
Return Loss and SWR

---
What is Group Delay?

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Basic Measurement Capability (S-Parameters)
Definition of Scattering Parameters
Measurement Concept
Flow Graphs
Cassini Measurement Hardware
Error Correction
Two Port Error Model

---
What is an S-Parameter?

---
Measuring 2 port s-Parameters

---
Second half of the S-Parameter
Measurement Process

---
Testing an Attenuator

---
S-Parameter Flow Graphs

---
One Port S-Parameter
Measurement Errors

---
Directivity (Directional) Errors

---
Reflection Tracking Errors

---
Test Source Mismatch Errors

---
Why we use Error Correction!

---
Expanded 2 Port Cals

---
Flow Diagram of 2 Port Errors

---
2 Port Calibration &
Vector Error Correction
Characterize Systematic Errors
Remove Systematic Errors from Measurement
Reflection Cal: Ms, D, Tr
Through Connection : Tt, ML
Isolation Measurement : C

---
4 Port Testset Config for
S-Parameter Measurements

---
Receiver S-Parameters Measurements

---
Forward or Unidirectional
Fastest method
6 Vector Measurements
No Source Switching
Allows "Hot" S22
Default RI Method

---
Forward Method
Measure a1, b1 and b2 terminated with known G1 and again with known G2.
Equations infer a2, and therefore s22
DUT

2

1
a
1
b
1
b
2

---
Forward Method

---
Conventional or Bi-directional
VNA Method
8 Measurements Required
Slower
Most Accurate
Switch Source to Output Side
No Hot S22

---
Conventional Method
Measure a1, b1, a2 and b2 ; Stimulated forward and reverse
DUT
b
a
a
b
1
1
2
2

---
Unidirectional or Bi-directional
Use Unidirectional for:
Speed
Hot S22
Use Bi-directional for:
S22, S12 accuracy

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Frequency Translation Device Measurement
Cassini Measurement Approach
Multi-Port S Parameter Detection Hardware 100MHz to 20 GHz
Requires one or more additional RF Source for DUT LO
Stimulus:
RF Stimulus Source: RF Input 100 MHz to 20 GHz
DUT LO Source: LO Input 10 MHz to 20 GHz
Maintain Constant IF Freq
Sweep RF and DUT LO Frequency
Requires 3 RF Sources to be at Different Frequencies
Measure: Pin(RF) and Pout(IF)
Calculate: Conversion Gain/Loss=Pout(IF)/Pin(RF)
Test Plan Optimizer Determines the Measurement Sequence

---
Frequency Translation Device Measurements

---
Frequency Translation Device Measurements

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
RF Power
Cassini Vector Corrected RF Power Measurements
Bandwidth Selectable Absolute RF Power (dBm, watts, etc.)
True RMS RF Power of Wide Band Digitally Modulated Signals
S Parameter Detection Hardware
Amplitude Only Measurement
IF Measurement
Wide Dynamic Range
Automatically Corrects for Signal Path Losses

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Spectral Purity (Harmonics or Spurs) Measurements
Cassini Measurement Approach
S Parameter Detection Hardware 100MHz to 20GHz
Relative RF Level Measurements
Stimulus
Single RF Tone at F1
User Specifies Harmonic No., N or Spur Frequency F2
Measure: Pout(F1) and Pout(F1xN) or Pout (F2)
Calculate: Pout(F1) - Pout(F1xN)) or Pout (F2)
Where Pout(F) = DUT Output Signal Level in dBm at Frequency F

---
Testset Configuration for Spectral Purity

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Intermodulation Distortion / TOI Measurements
Cassini Measurement Approach
S Parameter Detection Hardware
RF Level Measurements
Stimulus
Two Equal Amplitude RF Tones
Specify RF Level, Center Freq (F1) and Spacing
2nd RF Tone: F2 = F1-Spacing
Measure: 3rd Order Distortion Product Term
Pout(2xF1 - F2)
Calculate: IP3 / TOI in dBm:
Pout(F1) +[Pout(F1)-Pout(2xF1 - F2)] / 2

---
Testset Configuration for Intermodulation

---
RF Source Configuration for Intermodulation

---
Receiver Configuration for Intermodulation

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Wide Band Digitally Modulated Signal Measurements
Cassini Measurement Approach
S Parameter Detection Hardware
True RMS Noise Detector
Stimulus
Digital Modulation Signal Generator
Auxiliary Source Fast Power Control

---
Typical CDMA Signal
5MHz
5MHz
5MHz
5MHz
3.84MHz
3.84MHz
3.84MHz
3.84MHz
3.84MHz
3.84MHz
Delta F
ACLR = power in channel 3.84 MHz/Power in adjacent channel 3.84
MHx BW . Typical 55dB

---
Source 1 & 2 Combiner Module for Modulated Signals

---
Wide Band Modulation Receiver Measurements

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Digitally Modulated Signal Error Vector Measurements
Cassini Measurement Approach
Intrinsic Errors measured directly
I/Q DC Offsets and Mag & Phase errors for Mod/Demod
AM to AM and AM to PM for PA's
Stimulus
CW Signal Generator
Arbitrary Waveform Generator
Digital Modulation Signal Generator
"The Error Vector and Amplifier Distortion" - Published 1997 at the Wireless Communications Conference - Lucent / Bell Labs - Heutmaker
"WCDMA Transmit IC : Application case study" - Wireless Workshop - Published at the 2002 International Test Conference ( ITC) - Jointly by IBM and RI
" Edge PA EVM" - to be Published at the 2004 International Test Conference ( ITC) -

---
What is EVM?

---
This is EVM!

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Noise Figure Measurements
Definitions
What is Noise Figure
Y Factor Type Measurements
Second Stage Error Correction
Mismatch Effects
System Noise Figure Effects

---
What is RF Noise?

---
What is Nosie Figure?

---
S/N is too Difficult to Measure Directly - Additive Noise is Not

---
2 Port device NF Testing Model

---
2 Port NF Measurement Process
Y- Factor Technique

---
2nd Stage Error Effects

---
Gain of Device needed to extract
Noise Figure

---
Noise Source
External Calibration Standard

---
NF Uncertainty - Gain helps!

---
Single or Double Side Band Noise Figure?

---
4 Port Testset Config for
Noise Figure Measurements

---
Receiver Noise Figure Measurements

---
RF Measurements Outline
- Basic Measurements
- Phase and Magnitude
- Scattering Parameters
- Conversion Gain/Loss Measurements
- RF Power Measurements
- Spectral Purity (Harmonics) Measurements
- Intermod Distortion Measurements
- Digitally Modulated RF Signal Measurements
- Error Vector Magnitude (EVM)
- Noise Figure Measurements

---
Preview Next Chapter Developing and Running RI Test Plans
- Test Plan Concepts
- Test Plan Development Overview
- Creating a New Test Plan
- Test Plan Structure
- Building a Test
- Saving Device Data
- Viewing Tester Configuration
- Compiling & Running
- Viewing Test Results
- Setting Limits
- Release for Production
- Lab C

---
Questions?
Any Questions from this Chapter?