Cassini Maintenance and Diagnostics Training
Cassini
Maintenance Training
[Speaker Notes: Welcome to Roos Instruments’ short, Cassini Maintenance Training. This is also available as an intense 5 day seminar, combining class room training with hands-on experience. The Roos Cassini basics class is a prerequisite for this seminar. Also, it is highly recommended that the Cassini Applications class be completed before attempting this seminar.]

---
Cassini Maintenance Training
- System Introduction
- Routine Care & Maintenance
- Troubleshooting
- Wrap-up and Review

---
System Introduction
- Identification (Front/Back & Testhead)
- Block Diagrams (40 GHz)
- Power Supply, ON/OFF
- Emergency Off (EMO)
- Aux Rack
- Guru & System Controller Disaster Recovery
- Starting Cassini
- RIFL

---
Cassini 16 Front
Front

---
Cassini 16 Back & Testhead

---
System Block Diagram
40 GHz Cassini Block Diagram
See Appendix

---
PowerPoint Presentation
Cassini Device Power Diagram
See Appendix
[Speaker Notes: The DC and analog circuits are all in the bottom of the testhead. They are mostly independent circuits, consisting of 1-3 modules each. As such, the DC is normally rather easy to debug.]

---
Test System Power
Test System Power
- Always power down a module before working on it
  - TIMs are plug&play and can be removed without powering down the rack
  - Aux Rack sources are powered from the AC line
  - System Controller should be “Shut Down” via software prior to removal (watch for “blink animation”)
  - Always switch off Test Head +48V breaker switch on side of rack before opening Testhead (note: also powers down system controller)
- You don’t need to power down the entire test system.
- Remember the RIFL cable has power too (+48VDC)
Note: You do not need to shut down the entire test system for most service.
[Speaker Notes: Always power down a unit prior to working in it. Note that you do not need to power down the entire test system. In fact, we recommend that you not power down the entire test system to avoid extended warm-up times. Different units receive their power in different ways: The sources and hubs are powered by the test system’s main AC line through the UPS. To power down a source, unplug its AC main. Other components are powered by the RIFL. To power down one of these, such as the receiver, unplug the RIFL cable. Note that the RIFL cables may have 48VDC on them. Use caution to prevent shorting this DC. Fixtures should be removed prior to cycling power on the system. When a system is powered up, prior to the software taking control, it can come up in any state, including the possibility of voltage on 3 amp supplies. Removing the fixture ensures this momentary undefined state won’t damage the fixture, the DUT or the test system. Remember, there is a minimum warm up time of about 30 minutes when the test system is shut down completely. This is the case for 2 reasons; 1) The sources have ovenized quartz crystal oscillators in them, which go "cold" when power is turned off. 2) The temperature in the rack, particularly for the receiver, has to stabilize.]

---
Internal Breaker
- Test Head
  - Internal Test Head RIFL hubs
  - System Controller
- TIM Front – 8 TIMs farthest from the tower
- TIM Rear – 8 TIMs closest to the tower

---
System Power
Side Cover Removed
Cassini 16, Front Cover Removed
[Speaker Notes: The newest testers have a “Power Conditioning Sub-system” (PCS) instead of a UPS. The PCS performs the same function of conditioning the power. In addition, the PCS offers an emergency off button that shuts the tester off in case of emergency.]

---
Entire System Shut-down
Entire System Shut-down
- Shut-down the System
  Controller
  - Shut-down the operating
    system
- System Power
  - Turn the switch to OFF
  - Note that the power supply's
    startup circuit is still active.
- Main Breaker
  - Shuts off AC at the mains line
[Speaker Notes: Always remember to properly shut down the system controller's operating system before turning off the main switch.]

---
Emergency Off (EMO)
- Shuts off the entire tester in an emergency
  - To shut off the tester, press the button.
  - To re-start the tester, twist the button until it pops back out, then turn the System Power switch to 'Start'.

---
Aux Rack Shut Down/Turn On
- The Aux Rack is powered separately from the main tester
- The circuit breaker is also the EMO

---
A Note About Guru
Automatic Backup and Version Control

---
Guru Reliability
- All guru objects are tamper proof, hash verified and secure (no accidental overwrites)
- Complete validation and backups
- All programs are copied fresh and unzipped prior to execution (fresh)
- Latest version of every object retrieved instantly from Update Guru (no old versions)
- All versions (every save) is stored and backed up to Backup Guru, only removed by expiration policy. (never lose anything)

---
System Disaster Recovery
- If System Controller Fails, only data currently in RAM is lost, all saved guru objects (including test data) are always copied to the Backup Guru Server
- FAST & EASY – Factory replaced with ID already loaded, setup network and OS settings
- Login with “DefaultUser”, “Roos” and update Guru Address Book to point to Update/Backup Guru, Restart and GO!
  If available, Guru Server auto-discovery and GO!
- Optional, use System | Syncronize used to restore the latest version of all guru objects owned by the GuruID

---
Entire System Turn On
- Main Breaker is on.
- MO button is deactivated (popped out).
- Breakers on the side of the rack are on.
- FIXTURE SHOULD NOT BE ATTACHED
- Turn the front panel
  switch to 'On', then
  momentarily turn it
  to 'Start'.

---
Starting the Tester's Image
- The test system software (image) is a Guru object
- After the System Controller has booted, Login and Select 'Shortcuts' & choose the appropriate Shortcut
- The test system will startup
  - First it loads and compiles patches as defined by the Shortcut
  - Then it finds the hardware and identifies instruments
  - Then it assigns symbolic names to the instruments (for example, 'source1') from the tester definition

---
RIFL Connection
A Note About RIFL
- Contains:
  - 10 MHz reference to which everything is synchronized
    - System time base is in the RIFL master in the System Controller
  - +48VDC to power some modules
    - If it doesn’t have a plug, it’s powered from the RIFL
  - Serial data lines
  RIFL II looks like a network cable
  BUT ITS’NOT!
    - If you plug it into a network hub, smoke results.
    - RIFL cables must be RF shielded, most network cables are not.
RIFL:
Roos Instruments
Fast Link
RIFL II
(shielded
RJ45)
RIFL III
(Round)
[Speaker Notes: Units within the system communicate through the RIFL. RIFL is more than a data cable. RIFL is a custom interface. Within the RIFL interface is digital communications, +48VDC to power some of the units, and the system’s 10 MHz time base reference. Note that everything within the system is synchronized to the system’s 10 MHz time base reference through RIFL cables. It also contains 48VDC which is generated in the hub or Rack 40V Power Supply and is used to power some of the other modules (i.e receiver, source1/2 combiner module). DO NOT PLUG A RIFL CABLE INTO A NETWORK INTERFACE!! Doing so will definitely destroy the network hub and possibly destroy the system’s system hub. RIFL cables must be replaced with shielded cable, available from Roos Instruments. The shielding prevents radiated signals in the system. DO NOT use standard "Catagory 5" data cable. Not only can standard category 5 cable cause problems, it can cause very subtle problems that are very difficult to diagnose.]

---
Cassini Maintenance Training
- System Introduction
- Routine Care & Maintenance
- Troubleshooting
- Wrap-up and Review

---
Routine Care & Maintenance
- Daily, Periodic and Annual PMC Procedure
- Basic Care with PKZ, Coax
- Service Plans
- Transient Instruments & Cak Kits
- Calibration Process
- Inspecting & Special Cal Buttons
- Diagnostic Testplan Review

---
Regular (Daily) Maintenance
- Inspect Socket and Load Board
- If changing Fixtures, inspect all RF and DC connections between Fixture and TIMs
- Inspect Docking Hardware (wear)
- Check Handler POD RIFL cable
  (no kinks, damages)
- Clean with 90%+ Isopropyl Alcohol &
  Dry Compressed Air to remove desbris

---
Preventive Maintenance Procedure
Preventive Maintenance Procedure
- Inspect and clean all connectors
- Inspect all cables and connections in the rack
- Clean connectors if required
- Clean air filters if required
  (always turn off each instrument before servicing an air filter)
- Run diagnostics and verifies
[Speaker Notes: Refer to chapters 1 & 2 of the Maintenance/Diagnostics training manual for complete preventive maintenance procedures.]

---
A Note about
Cleaning PKZ Connectors
- Spray an appropriate cleaner into the connector
- Dry with clean dry air or CO2
See “Cassini Fixture Care and Maint”
http://roos.com/docs/RBES-84B49L?Open

---
Handling coax cables
Handling coax cables
- Always turn the coupling nut, do not spin the center conductor
- Do not twist the cable
- On rear panel cables,
  do not try to put a
  wrench on the gold
  trim cap
[Speaker Notes: Many people, even those who have worked with RF, don’t know how to properly connect a coax connection. The key to any connecting any coax connector is to remember that the center conductor should not spin. The center conductors usually have a thin gold plate on them. If you spin the connector excessively, it can wear off the gold plate, which allows the connector to oxidize and causes poor connections. Note that, on the back of the Roos test systems, we use right-angle connectors. These connectors have a decorative gold cap on them that looks like it can take a wrench. IT CAN NOT. It is only glued on, and will come off if you try to use a wrench on it.]

---
Access to Service Plans
Access to Service Plans
- Virtually all service activity is performed through the system's ‘Configuration’ window
  - To open the configuration window, from the Cassini application window, select the ‘System’ button, then ‘Tester’
- If you don’t know what it does, DON'T TOUCH IT.

---
Types of Service Plans
Types of Service Plans
- Calibrate
  - Changes cal factors
  - Don’t run if you don’t know what you’re doing
- Validate
  - Validates that the cal ran properly
  - Only valid immediately after cal
- Verify
  - Verifies that a calibrated system is working
  - Run any time
- Diagnose
  - Verifies that Hardware is working
  - Theoretically does not require calibration
[Speaker Notes: We break our service plans into 4 categories: Calibrate, Validate, Verify and Diagnose. Calibrate: Precisely as the name implies, these plans perform calibrations on the system. Measurements are taken using calibration standards and cal factors are stored in the test system. Validate: Validates check the results of the calibration to very tight standards (tighter than spec, to ensure that the test system won’t drift out of spec over time). They are only valid when run immediately after calibration and are only accessed as part of the calibration process. Verify: Verifies check that a properly calibrated test system is working to spec. Verifies can be run any time to check the performance of the test system. Note that verifies check the quality of the calibration as well as the quality of the hardware. Diagnose: Diagnose plans check that the hardware is working. Diagnose plans are theoretically independent of calibration, however no test plan can be fully independent of calibration. Note that sometimes diagnose plans can fail if the calibration is very far off. Diagnose plans can be run at any time and should be run periodically, along with verifies, to confirm proper operation of the test system.]

---
Recommended PMC Schedule
Periodic & Yearly PMC Sched.
- Calibration at 1 year intervals
  - Includes diagnose/verify
- Diagnose/Verify as required
  - Some customers perform at 1 month intervals
  - Some customers perform only if a problem is noted (recommended)
[Speaker Notes: Preventative Maintenance and Calibration (PMC) consists of 2 parts: Diagnose and Verify confirms that the test system is operating properly. Calibration re-adjusts the calibration factors to account for changes over time. Calibration should be performed every 6 months. Normally, RI is contracted to perform the cals. At that time we check the entire tester for potential problems. If you choose to do calibrations yourself, be sure to check for and fix any problems before attempting the cal. Diagnostics and verifies should be run as often as you feel is necessary. Some customers run them as often as once/month. Some don’t run them at all, leaving it to RI to run at the 6 month calibration.]

---
Activating a Cal Fixture
Activating a Cal Fixture
- Cal fixtures should activate automatically when a 'System' 'Check' is performed
- To activate manually, from the system's 'Configuration' window, select 'Instrument' 'Add Fixture Def'
- If another fixture is active, the system cals will be wrong, because of that fixture’s cal data.
  - Cals are just test plans with a couple of special buttons for resetting/saving the cal data.
  - A DUT interface can also affect the cal if the DUT interface contains cal data.
[Speaker Notes: Always activate a calibration fixture when performing service. If a Calibration Fixture is not activated, the test system may use the calibration factors from the active fixture, that can cause erroneous test results. Note that some service plans will not run unless the Calibration Fixture is activated. However, some will, so always be sure to manually activate the Calibration fixture.]

---
Transient Instruments
- Instruments such as Power Meter and Cal Kit are “transient” (meaning they are not always connected to the system)
- Transient instruments are listed in the system's Configuration window
- Always remove a transient instrument before disconnecting
  - To remove any instrument, Right click on the instrument in the system's Configuration, then “Remove”
[Speaker Notes: Instruments such as the power meter and the cal kit are ‘transient’ instruments. That means the test system does not assume they are always connected. You must tell the test system when a transient instrument is connected. Transient instruments are found in the ‘System’ ‘Tester’ panel. To activate a transient instrument, be sure it is physically connected, then Select 'Instrument' 'Add Instrument” Be sure to de-activate a transient instrument before physically disconnecting it. Otherwise, you might get error messages the next time the tester is used. Note that the same Giga-tronics power meter name is used for all power meters in Giga-tronics’ 854X family.]

---
Activating a Cal Instrument
Activating a Transient Instrument
- Activating a Transient Instrument is done in the Configuration Window
- Select “Instrument” “Add Inst”
  - The tester will ask you for a name. Generally, accept the default name.
  - For GPIB instruments, the tester will ask for a GPIB address.
    - Roos always puts power meters at address 5.
- Although the cal kit is considered an instrument, it is activated differently

---
A Word About Cal Kits
A Note About Cal Kit Maintenance
- Cal Kit:
  - DMM: Standard for DC (Voltage & Current)
  - Power Meter: Standard for RF Power
  - Open/Short/Load (OSL): Standard for RF impedance
  - Noise Standard: Standard for Noise Figure
  - Cal/Diag/Engineering Fixture
- Cal kit components must be calibrated
  - Power meter and DMM: 1 year
  - OSL and Noise standard: 2 years
[Speaker Notes: When calibrating a test system, a cal kit is used. The cal kit provides the traceable standards necessary to fully cal a test system. The calibration interval for an CASSINI is 12 months. Note that the cal kit components must be calibrated from a standards-traceable cal facility. Roos recommends that, whenever practical, the cal components be sent to their manufacturer for cal. We recommend most cal components be re-calibrated once per year. Cal Standards other special cal devices can be calibrated once every 2 years.]

---
Updating a Cal Kit Definition
- After calibration, cal kit coefficients may need to be changed
  - Configuration window, ‘Instrument’ ‘Calibration’ ‘Inspect’
  - Double-click the appropriate element.
  - Entries default to linear.
    For log, click the ‘lin’ button
[Speaker Notes: Always power down a unit prior to working in it. Note that you do not need to power down the entire test system. In fact, we recommend that you not power down the entire test system to avoid extended warm-up times. Different units receive their power in different ways: The sources and hubs are powered by the test system’s main AC line through the UPS. To power down a source, unplug its AC main. Other components are powered by the RIFL. To power down one of these, such as the receiver, unplug the RIFL cable. Note that the RIFL cables may have 48VDC on them. Use caution to prevent shorting this DC. Fixtures should be removed prior to cycling power on the system. When a system is powered up, prior to the software taking control, it can come up in any state, including the possibility of voltage on 3 amp supplies. Removing the fixture ensures this momentary undefined state won’t damage the fixture, the DUT or the test system. Remember, there is a minimum warm up time of about 30 minutes when the test system is shut down completely. This is the case for 2 reasons; 1) The sources have ovenized quartz crystal oscillators in them, which go "cold" when power is turned off. 2) The temperature in the rack, particularly for the receiver, has to stabilize.]

---
Verify that the System Knows the Cal Kit Information
Verify that the System Knows the Cal Kit Definition
- The system must have the proper Cal Kit definition
- Coefficients are stored in the Cal Kit definition
  - OSL coefficients, particularly the open's fringing capacitance
  - Noise standard's ENR table
- Cal kits are activated thru the Configuration window
  - “Instrument...” “Add CalKit”
[Speaker Notes: To verify the system knows the cal kit information, refer to the "Adding a Cal Kit Definition to a Cassini System" document, available on the Roos Instruments web site, for instructions. Also, if the cal kit you are using is different from the one listed in the instruments panel, or has recently been calibrated, you will need to create a new cal kit definition, or update the existing cal kit definition. Instructions for this are also available in the ‘Adding a Cal Kit Definition to a Cassini System" document on the web site. Step 1: Activate Correct Cal Kit Select System | Tester Highlight Testhead, select “Instruments”, “Calibration”, “Calibrate” Select “Options”, “Select Cal Kit”, highlight and select the appropriate Cal Kit Step 2: Go To Configuration Panel, Highlight Cal Kit, select “Instruments”, “Calibration”, “Inspect” Step 3: Inspect and check that it has correct Data]

---
Calibration Process
Calibration Process Overview
- Open the appropriate cal list from the configuration window
- Activate the appropriate instruments (cal kit and power meter) and activate a cal fixture definition.
- Perform Diagnostics
- Calibrate the test system
- Perform Diagnostics again
- Notes:
  - Calibration requires User in Maintenance Group
  - Cals should be run a few at a time
  - Remember, you have the power to damage the test system!!
[Speaker Notes: Always run diagnostics and verifies before calibration and after calibration. When calibrating a system, note that the system does not save the cal information until all the selected cals are run. Therefore, run the cals a few at a time. This prevents having to re-run a lot of cals if something goes wrong.]

---
Opening a Calibration Executive
- From the Configuration window
  - ‘System’ Calibration Calibrate
  - Choose the cal list

---
A Note About RF Cals
- DC and Analog are 'self contained'. There is no interaction with other TIMs
- RF interacts due to mismatch ripple and cable length
  - Individual TIMs are calibrated one time, at the factory
  - The 'testhead' cals serve as 'RF System' cals
    - Compensate for TIM-TIM interaction
    - Some cals can be affected if a source, combiner, or receiver is changed.

---
Show All Results
Show All Results
‘Options’ ‘Show All Results’
Always select “show all results” when taking data to send to RI
[Speaker Notes: When running any service plan, you have the choice to show fails only, or to show all results. If you are sending data to RI, always select ‘Show All Results’]

---
Cal Plan Summary
Cal Plan Summary
- DC Cals
  - Vmeasure
  - VCC
  - Parametric Meas/Supply
  - VI
  - Static Digital
- Waveform Cals
  - Digitizer (Oscilloscope)
  - Waveform Generator
  - Sinegen
- RF cals
  - Receive Attenuator
  - RF paths
  - EA (Error Adapter)
  - Direct Receive
  - Noise Cals
    - Receive
    - Source
[Speaker Notes: The cal plans are organized within each Instrument in the order they should be run. We’ll review each of the cal plans with brief descriptions of each one. Special notes about cal plans: - Cal plans do not pass or fail. They only report ‘done’. To confirm that the cal performed properly, a validate or verify is run after the cal. - The cal plans contain buttons that reset cal factors. Note that the cal factors are reset when the cal plan is compiled, not when it is run. This means if you abort a cal plan in the middle of its compile, before it has run, the cal factors may have been reset to default values. - Your cal plan list might be different, depending on your option mix, and the vintage of your test system and which TIMs are installed. - Note that you need to be logged on as a super-user to run cals. Remember, as a super-user, you have the ability to render the system totally inoperable. DO NOT RUN CALS, or make any changes to the system, unless you know what you’re doing.]

---
40 GHz Cal Plan Summary
- Source Slope
  - Gets source close at one power
  - 'Iterative' cal, may need to be repeated
- Port 1 Rec Freq Response
- Source Linearity
- Source Freq Response
- EA Error Adapter)
- Port 2 Rec Freq Response
[Speaker Notes: The cal plans are organized within each Instrument in the order they should be run. We’ll review each of the cal plans with brief descriptions of each one. Special notes about cal plans: - Cal plans do not pass or fail. They only report ‘done’. To confirm that the cal performed properly, a validate or verify is run after the cal. - The cal plans contain buttons that reset cal factors. Note that the cal factors are reset when the cal plan is compiled, not when it is run. This means if you abort a cal plan in the middle of its compile, before it has run, the cal factors may have been reset to default values. - Your cal plan list might be different, depending on your option mix, and the vintage of your test system and which TIMs are installed. - Note that you need to be logged on as a super-user to run cals. Remember, as a super-user, you have the ability to render the system totally inoperable. DO NOT RUN CALS, or make any changes to the system, unless you know what you’re doing.]

---
Inspecting Cal Data
Inspecting Cal Data
- ‘System’ Button ‘Tester’
  Select desired instrument
  ‘Instrument’ ‘Calibration’ ‘Inspect’
- View data tabular or graph
- Never edit cal data unless specifically Instructed by RI
  - You can get the test system into a state where it won’t even start up
[Speaker Notes: A BAD CAL CAUSES BAD DATA!! Often a problem is found when a cal is run, then the test system fails the validate. Then the problem is difficult to debug because, not only do you have a bad piece of hardware, you’ve now got bad cal data. This can usually be avoided by always running the diagnostics and verifies before starting a cal, and if a cal does not complete, restore from a backup of the tester with Guru Browser! You can view the cal data to see if it looks reasonable as follows: In the Cassini Window, select 'System' then 'Tester', select the desired instrument (for example 'testhead'), select 'Instrument', 'Calibration', 'Inspect'.]

---
Test Plans Special Calibration Buttons
Special Calibration Buttons
- Special buttons
  - Reset cal factors
  - Write to cal factors
  - Perform special calculations for calibration
- Not available through normal editor
- Do not copy or modify special buttons
[Speaker Notes: Special Calibration Buttons are just that “Special”. Cal factors can be written to, and reset through these buttons. They also perform certain calculations for cal. Bottom line in regards to special cal buttons, DO NOT TOUCH!!]

---
Diagnostic Plan Review
Diagnostic Plans
Checks the noise source and the noise receive path. Note that these can fail if receive
path has too much loss
Checks the internal noise reference
Measures receive path IP3. Note: Most failures on this test are, in fact, path loss problems, not IP3 problems
Checks the bandwidth and rejection of filters in the receiver
Checks the receiver's detectors and IF gains

---
Diagnostic Plan Review
Diagnostic Plans
Checks source/receive RF relays in test head (11, 12, 21, 22)
Checks the receive attenuator’s relative attenuation
(receive attenuator is in the test head.)
Checks that signal is getting to the ports and that they are a good match
[Speaker Notes: All Roos service plans are the same as the test plans that you write, except for the few special calibration buttons. If you can read a test plan, it’s normally easy to see exactly what a diagnose plan is doing.]

---
Diagnostic Plan Review
Diagnostic Plans
Checks the source1/2 combiner's switches and attenuators
Checks the sources produce RF properly
Checks the coaxial switches in the test head for through loss and match
[Speaker Notes: Note that some of the Diagnose Switch tests set the test system up in strange ways, to measure match and through loss of the switches.]

---
Cassini Maintenance Training
- System Introduction
- Routine Care & Maintenance
- Troubleshooting
- Wrap-up and Review

---
Troubleshooting
- Breakpoints & Control
- Limits, Graphs
- Disabling Tests
- Basic, Normal Troubleshooting Tree
- Common Senario (Example)
- Source Basics, Block Diagram
- Common Source Failures
- Troubleshooting Hints
- Find Phase Lock Errors with Averaging

---
Test Plans- Setting Break Points
Test Plans - Setting Break Points
- ‘Options’
  ‘Set Break Point’
  - Changes color at the break point
  - Can only break on a test, not on defaults etc.
  - Test plan will stop after setting up the hardware and taking the measurements
  - Breakpoint stops after the post-measure
  - Only 1 break point per test plan
  - You can continue, single-step, abort, or open the control panel
Note: Do not do a system start-up or check while at a break point
[Speaker Notes: A breakpoint in a test plan allows you to stop in the middle of the plan. Breakpoints are useful for debugging test plans as well as debugging test system problems. To set a breakpoint, in a test plan, select ‘options’, then ‘set breakpoint’. Within the test plan, the test where the breakpoint is set will turn ‘Blue’. If you select some other test to view, the test with the breakpoint will turn yellow to indicate the breakpoint is still active. In the example above, the breakpoint on ‘High Band’ is yellow because you’re looking at ‘Low Band’. Note: Do not do a system startup while at a breakpoint because it leaves the system in an undefined state. If you happen to do a start-up while at a breakpoint, wait until the startup process finishes, then select ‘Abort’ on the breakpoint panel. You’ll probably get a ‘walkback*’ error. If so, select ‘Abort’. Then do a system start-up again. *A walkback error is an error message generated by the system’s operating system. Unusual errors that confuse the operating system and are not specifically trapped by the software will result in walkback errors. When you get a walkback error, always select ‘Abort’, then fix the condition that caused the error.]

---
Test Plans - Control Panel
Test Plans - Control Panel
‘Options’ ‘Control’
        or
From Breakpoint
  - If reached from a breakpoint, reflects the settings at the breakpoint. Remember that the testhead is ‘hot’ at this point.
  - Manually control system instruments
[Speaker Notes: The ‘Control Panel’ allows you to manually control system instruments. As in this example, source 1 has been selected, then ‘control panel’. What you are looking at is the present state of source 1. If the control panel is opened while the test is at a breakpoint, the control panel will reflect the state of the system at the breakpoint. * Remember, the testhead is “HOT” at this point. When using the control panel, be careful not to apply excess voltage or RF power.]

---
Test Plans Disabling Parts of Tests
Disabling Parts of Tests
  - Disable parts of tests to increase speed or decrease confusion
  - Programmer message window will always say ‘Fail’ if any test or test section
    is disabled
  - ‘Edit’ ‘Disable/Enable’
      ‘Selected’
           or
     ‘Disable All’
    then re-enable selected
[Speaker Notes: There are a couple of reasons you might need to disable parts of a test during debug. One is to increase the speed of a testplan. Disabling one of the many tests or test sections of the testplan will make it run faster since there are now fewer tests being run. This is particularly useful when doing a ‘rattle’ test. Another reason is to decrease confusion. If you are troubleshooting a problem, by disabling a test or test section, it will decrease the amount of data that you have to sift through while searching for a solution. To disable a test, first click on the test you want disabled. Then, at the top of the testplan window, select ‘edit’, ‘disable/enable selected’. This will disable the slected test. If you want to disable more than one test in a section, repeat the process, then select ‘disable all’. This will disable all the tests in that section.]

---
Test Plans Aborting a Test Plan
Aborting a Test Plan
- Validate, Verify, Diagnose can be aborted if necessary
  - ‘Ctrl’ ‘Break’
  - System will show a walkback error. Select ‘Abort’
  - From Cassini's ‘System’ menu, select ‘Reset’
    - Always do a system reset after aborting a test plan
  - If a plan is to be re-run, select ‘Debug’ ‘Clear Compile’
- Cals can be aborted, but only if absolutely needed
  - Cal data will be invalid
Note: If a validate fails, the cal data is not saved. It can be manually saved by RMBC on the appropriate instrument and selecting ‘Save’.
[Speaker Notes: Sometimes it may be necessary to abort a testplan. The reasons can vary, from missing a step specified in a condition statement, cal standards put in the wrong port, etc. All testplans can be aborted. Cal plans should only be aborted if absolutely necessary. To abort a testplan, do a ‘control’ ‘break’. (hold ‘control key down, then hit the’break’ key). A ‘Walkback’ error will appear. Select ‘abort’. Then, in the ‘Programmer Message Window’ select ‘System’ ‘Reset’. (always do a system reset after aborting a testplan). Most calibration plans reset the cal data when they compile, then insert new cal data when they run. If you abort a cal plan, or compile it but don’t run it, you probably have lost the old cal data. Normally when you run a group of cals/validates through the calibration executive, when they are finished you will see a message that says ‘saving cal data’. However, if the ‘validate’ fails, the cal data is not saved. If you want to save the cal data, it can be saved manually by doing an ‘RMBC’ on the active tester, then selecting ‘Save’. If either of the above has occurred and the cal data has not been saved, you can still get the old cal data back by de-activating the tester (RMBC on the tester icon and select ‘de-activate’). When it asks if you want to save cal data, select ‘No’. Then, it doesn’t save the new data and the original cal is restored.]

---
Test Plans Finding Limits
Test Plans - Finding Limits
- Limits are helpful for seeing how close the instrument is to failing.
- Limits are set on data save buttons
- To check limits:
  - ‘Limits’ ‘Select’ ‘SystemCheck’
    - All limits for cal/service plans are called SystemCheck
RMBC on data save button
‘Single Value Limit’ (same as customer test plans)
NEVER CHANGE A SERVICE TEST PLAN!!!
[Speaker Notes: All of the validate/verify/diagnose plans have limits. If the measured data exceeds the limit, the test fails. Sometimes a diagnose/verify/validate will fail by a small amount. You will need to know the limit of the test, to know how far off the failure is. To check the limits of a test plan, open the test plan, re-name it to ‘junk’, and compile it. Select 'limits', 'select', then ‘SystemCheck'. The limits for the calibration & service plans are all titled ‘SystemCheck'. Right Mouse Button Click (RMBC) on the data save button, select, 'Single Value Limit'. This will display the limits for that measurement. To find the limits within a test plan, you first must find the data save button(s). NEVER CHANGE A TEST PLAN! Remember that you are a super-user and can change things that will cause the test system to malfunction. As noted in the slide, always re-name a test plan to ‘junk’ before working with it. When re-naming a test plan, if ‘junk’ already exists, it’s OK to write over it.]

---
Test Plans - Finding Limits
Test Plans - Finding Limits
[Speaker Notes: This slide shows the limits for Vcc1 in the ‘VCC 123456 Validation CF2’ test plan. As before, Right Mouse Button Click (RMBC) on the data save button, select, ‘Single Value Limit’. The range for Vcc1 in this test plan is a minimum of -0.02V (-20 mV), and a maximum of 0.02V (+20 mV).]

---
Test Plans Viewing Data as Graphs
Test Plans - Viewing Data as Graphs
- Left click on Data Save or Local Variable button
- Select graph type
  - Rectangular for arrayed data (freq. Response data, etc.)
  - Moving strip chart for CW data
    - Moving strip chart can handle up to 500 measurements
- If local variable, will need to format graph
  - Sort by ‘Index’, X value sort by ‘Index’
[Speaker Notes: Normally, it’s easier to interpret data as a graph, rather than a bunch of numbers. As such, one of the more powerful troubleshooting tools is the test system’s ability to graph data. Many graph types are possible, but the 2 that are used 90% of the time are the rectangular graph and the moving strip chart. If a test does multiple data points in one data save button (for example, if the test does measurements at multiple frequencies), then the rectangular graph is used. If the test saves one point, then the test can be repeated, and the moving strip chart is used to check repeatability. The moving strip chart is particularly useful for ‘rattle’ tests, where you’re looking for things that change with mechanical movement, such as loose cables. Note that you can graph local variables, as well as data saves. However, if you graph a local variable, the data will be pre-sorted in a manner that you probably don’t want. To view the data ‘as taken’ on a local variable, you need to tell the graph to sort by ‘Index’, and to sort the X value by ‘Index’]

---
Test Plans - Viewing Data as Graphs
Does this graph show good data or bad data?
[Speaker Notes: A word of caution about ‘Auto’ scale. It’s possible to take perfectly good data and make it look bad by selecting the ‘Auto’ button. This button will increase the scale until the data fills the screen. This graph is essentially the same as the previous graph. Note, however, that the Y SCALE is 0.001 dB/division. Always be aware of the scale when viewing data as a graph.]

---
Troubleshooting Basics
Troubleshooting Basics
1. Is the problem in the source path or receive path?
2. Does the problem
  - Change with respect to frequency?
  - Change with respect to power?
1. Murphy’s corollary: There is always more than one problem
2. One problem can cause multiple failures
  - Be sure you’re looking for the right problem
  - Usually run all diagnostics
[Speaker Notes: The above slide shows the basic thoughts to always keep in mind when troubleshooting the Cassini Test System. The test system uses the sources to test the receiver, and visa-versa. When a test fails, you can immediately cut the problem in half by checking if the problem is in the source path or the receive path. Also, does the problem change with respect to frequency (frequency response problem) or power (linearity problem)? Always remember, more often than not, “Murphy’s Corollary” applies. “There is always more than one problem”. The inverse of that can also apply: sometimes one problem can cause multiple failure symptoms. This is why you want to run all the diagnostics, to find patterns in the failure symptoms that point to a specific problem.]

---
‘Normal’ Troubleshooting Tree
‘Normal’ Troubleshooting Tree
- Run Diagnostics (all of them)
- Review failed data
- Look for patterns (only 1 port, direct receive path, etc.)
- Pull up test plan
- Re-name to ‘Junk’
- Set breakpoint
- Run to breakpoint
- Open Controller
- Manipulate settings through controller
- Begin ‘halving’ signal path
- Isolate bad component
[Speaker Notes: Unfortunately, for RF, it’s not practical to insert test points wherever we feel like. So, to diagnose RF problems, you normally have to run all of the tests, then look for the patterns in the failures. For example, if a problem occurs at RF3, but not RF2, 6, or 7, then the problem must be in some area that’s unique to RF3. If the problem occurs at RF2, 3, 6, and 7, then the problem must be with something that’s common to all those ports. Using this process of elimination, you can usually diagnose the problem to a couple of possibilities before you connect any test equipment. Once you’ve done the heavy thinking to eliminate what it can’t be, then you can check for the remainder by manipulating a test plan to stop at specific points, and checking where the signal goes bad. Remember, ALWAYS RENAME TEST PLANS TO ‘JUNK’ BEFORE YOU MAKE ANY CHANGES. You are a super-user. As such, you have the power to render the test system totally inoperable. An easy way to ‘accidentally’ do this is to modify a test plan then save it.]

---
Common Scenario
Common Scenario
- During Cal, a validate fails
- You troubleshoot the problem to a bad switch
- You replace the switch
- You run diagnostics, and now diagnostics fail
What’s wrong?

---
Synthesizer Block Diagram
Source Block Diagram
YIG Osc.
Switched Filter
RF
Out
Coarse
Loop
ALC
Down Conv.
YIG
Loop
Fine Loop
Ref Loop
OCXO
System
Time
Base

Det.
DDS

Det.

Det.
Reference loop: locks the source to the RIFL time base
Coarse Loop: coarse steps, L.O. for YIG loop
YIG Loop: Main loop, locks YIG oscillator
Fine Loop: 1 Hz steps, reference for YIG loop
[Speaker Notes: The synthesizers contain as much circuitry as everything else in the test system combined. The fact that they are complex and that there are 4 or 5 of them in a system means the synthesizers are the most common failure item in a system. We don’t expect you to become experts in the design of microwave synthesizers. However, a basic understanding of how the sources work and how they are used in the system will help you troubleshoot source problems and perhaps even repair sources on-site. A synthesizer simultaneously achieves high resolution and low noise through a complex arrangement of phase lock loops. There are 4 loops in a synthesizer, as shown. The YIG (Yttrium Iron Garnet) oscillator produces the actual output. It is able to be tuned from 2 GHz to 20 GHz. For frequencies below 2 GHz, the YIG oscillator is mixed with a fixed frequency oscillator in the Down-Converter to produce 10 MHz to 2 GHz.]

---
Source Common Failures
  - A3, reference loop
    - Certain vintages
    - Time base ‘shifts’ after 1-3 years
  - Down Converter
    - Dead
    - Phase lock (frequency accuracy) problems
    - ALC (power control) problems
[Speaker Notes: The most common failures are shown. What symptoms would you note for each of these failures? ___________________________________________________________________ ___________________________________________________________________ ___________________________________________________________________]

---
Source Basics
Startup Source Tests
  - Check all loops
  - Check ALC range
  - Single frequency test
Self Test (Control Panel)
  - Check all loops
  - Check ALC range
  - Multiple frequency test
[Speaker Notes: There are many self-tests that help to diagnose source problems. The nature of synthesized sources is such that it is relatively easy to confirm all the loops are properly locked. The software gives the sources a quick test whenever a system startup is performed. A more comprehensive test can be performed by running the self test in the control panel. To run the self test, open the control panel, highlight the appropriate source, and select ‘Instrument’ ‘Self Test’. The test system will run the self test and show the results in the programmer’s message window.]

---
Source Troubleshooting Hints
Source Troubleshooting Hints
Problem only below 2 GHz
  - Likely down converter
- Different averages can help diagnose lock problems
- Power meter can eliminate confusion from receiver issues (remember Murphy’s corollary)
- Power correct can be turned off (but not all corrections are removed)
[Speaker Notes: With the internal self-tests and your knowledge of sources, source problems are usually straight-forward to find. One exception can be subtle phase lock problems, where the self test can’t detect the problem.]

---
Why Averaging can Find Phase Lock Errors
Why Averaging can Find Phase Lock Errors
Q
Q
Amplitude
Phase Angle
I
I
Effect of averaging on a phase
locked (synchronous) signal
Effect of averaging on a non-phase
locked (asynchronous) signal
Averaged
Measurement
Averaged
Measurement
[Speaker Notes: Changing averages can find source phase lock errors. This is because the average is a vector average. If the source is not phase locked, its phase will be randomly changing with respect to the receiver. When averaged, the result will be a smaller amplitude. In the diagnostics for each source, there is a phase lock test that performs this function.]

---
Cassini Maintenance Training
- System Introduction
- Routine Care & Maintenance
- Troubleshooting
- Wrap-up and Review

---
Wrap-up and Review
- Exchanging TIMs
- Sending Data to RI
- Guru Object Class Definitions
- How to Get Information
- How to Get Help from RI

---
Exchanging TIMs
- TIMs can be hot-swapped (DO NOT power down the system)
- Pull top to prepare, hold tight and pull bottom to release TIM
- Get RMA from support@roos.com
- Ship in double box
  - Inner origional package (carboard & plastic)
  - Outer Package (at least 26" x 20" x 8")
  - At least 3” of packing material on all sides

---
Sending Data to Roos
Sending Data to Roos
- Often we can be of more help if we can see the test data or the calibration data.
- To send the cal data, use Guru Browser to export a “.gzp” (Guru Export) of the instrument's cal data.
- Filter Guru Objects with the following keys:
  - ri.sys.ObjClass = RiInstrumentCal
  - ri.hw.Model = <Instrument Model Number>
  - Right click, select “export” Guru Format
- To send results from Diag testplans, use “Results | Save to Guru” and export “RiEquipDiagLog” object class.
- Or the results can be saved as text (.txt) file.

---
Important Tester
Guru Object Classes
- RiBaseInstDef – Instrument Feature Definitions
- RiCalibrationKit – Calibration Offsets
- RiCalList – Calibration Test Plans
- RiEquipDiagLog – Diagnostic Logs
- RiFixtureDef – Fixture Definitions
- RiFixtureCal – Fixture Calibrations
- RiInstrumentDef – Instrument System Drivers
- RiInstrumentCal – Instrument Calibrations
- RiPatch – Cassini System Software Update
- RiShortcut – Cassini System Software Management
- RiTesterDef – Cassini Instrument Collection Definition
- RiTestplan - Testplan

---
How to Obtain Information
How to Obtain More Info
- Roos Instruments Web Site
  - Wealth of Support and Applications information
  - www.roos.com
  - ‘Customers Only’ for the good stuff.
  - Log-in required (register online)
[Speaker Notes: The Roos Instruments web site is a valuable source of information. To get to the good stuff, you need to log in with a password. If you don’t have one, email ‘support@roos.com’ to request one. When you send information to Roos, always send it to ‘support’. This guarantees that it will always get attention. If you email a specific person and they are away, you may wait days for a response.]

---
How to Get Help from RI
- Roos Instruments
  - +1 408-748-8589
    Mon-Fri: 8-7pm PST (-8 GMT)
- Service contracts offer phone/email support and fast turn-around for parts
  - Check for the local support phone number
- Email: support@roos.com
  - Routes the email to several support people so it never gets missed.

---
Cassini Maintenance Training
- Questions?
- Hands On