Fixture Calibration
Calibration Types
Direct
  - Uses a primary NIST traceable standard (OSL, Power Meter, O-scope).
Implied
  - Uses a secondary standard (one calibrated resource to calibrate another).

---
Fixture Calibration
Socket
DUT Interface Board (DIB)
Top Plate Assembly
Top Plate & Carrier Board
Modules
Modules
Calibrate to here and
calibrate DIB separately
Calibrate fixture to here
and treat as a single Cal
OR
Bottom Plate Assembly
Power VIs
RF
AC
GND
DC
Test Head

---
Cal Plan Creation
Using Templates
1. Go to main environment window
    and pull down the “Test” menu.
2. Choose “Plans....”.
3. When in the Test Plans window
    change directory to: RIAPPS\testplan\Caltests\Fixture\Standard Templates
4. From the listed templates choose
    the one that most represents the
    port you want to calibrate.
5. Copy it to a name that represents
     your device. Ex. RF000_RF2

---
Student Fixture Block Diagram
RF000
DutP25&26
DP1
Clk
VccPA
DutP30,31,&32
Data
Stb
PA out
FHSD
DutRF1
DutRF4
PA in
RF3
Q in
DutP2
RF6
Mod out
DutRF5
+
90
WF2
DutRF2
I in
DModin
DutRF6
RF7
DutP1
WF6 & RF4
I out
DutP3
WF3
WF7 & RF4
I Q out
DutP4
RF2
DutRF3
90
+

---
Editing the Port Template
RF000 Fixture
6. Open the new test plan in the edit mode. Also make sure that the fixture file is active.
7. Start in the Global Defaults and edit the RF path button.
RF Path to Calibrate

---
Editing the Port Template
8. Go to the “Open” test section and redefine the path to be calibrated.
RMBC then select
from list

---
Editing the Port Cal Plan
9. Go to the “Calc” test section and redefine the “Cal Data” data save button.
Select Cal Data Name
RMBC then select
CalVarName

---
Editing the Port Cal Plan
10. Go to the “Global Defaults” section and adjust the frequency span, frequency steps, power, averages, and IF gain to suit the specific application.
Power Level
Frequency Span and Steps
Averages
IF Gain

---
On Power Meter Cals
May Need to Reset
Data Format

---
Other Things to Look For
Look for warnings in the main message window
and adjust the IF Gain and RF source power
as needed.
Edit the Operator Pause as needed.

---
Checking the Port Cal Data
1. Go to “Fixtures” from the pull down menu in the main window.
1. Activate
2. RMBC
3. Calibration Inspect
Path Definition
Cal Data (default)

---
Viewing the Port Cal Data
2. High light the path definition and with a RMBC choose “View”.
View
Cal Data
Path Definition
3. Choose “Rectangular from the viewer.
Select

---
Viewing the Port Cal Data
3. Adjust “Y Parameter” to data type and “Y Format” to data type.
Auto Center
Y Parameter
Auto Scale
Y Format
Memo Field
4. Check for proper frequency range, loss level, and/or power level to make sure it represents the application requirements.

---
Good Data / Bad Data
Nice monotonically sloping data
RF3 to DUTRF thru path.
No series devices.
OSL calibration.
48MHz – 2GHz DUT test band.
Is this good cal data?
1.3 GHz
1. Fixture not calibrated to 48MHz.
2. Tester not calibrated to 48MHz.

---
Why Look at Cal Data?
RF3 to DUTRF thru path.
No series devices.
OSL calibration.
48MHz – 2GHz DUT test band.
Customer complained of 2DB offset at 1.3GHz between pre an post calibration.
1.3 GHz

---
Remember
Calibration plans are written from templates, copies of other cal plans, or from scratch and as such are prone to error.
Calibration test plans must be verified good and debugged to represent the application specifics (frequency range, power, calibration type, etc...).
It is absolutely necessary to look at the calibration data. NEVER assume it is correct.