Intro
Roos Instruments, Inc.
CASSINI
RF ATE System
Basic Operations
Training Seminar

---
Seminar Outline
Cassini Basic Seminar Outline
- Operation and Troubleshooting
- System Administration and Maintenance
- Basic test Plan Concepts
- Science of RF Measurement
- Device Definitions
- Example Application Development
- Test Fixture and Device Interface Design
- Test Design & Best Practice Test Optimization
- Application User Guides

---
Outline1
Chapter 1 Outline
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Guru Log-on, Applications
- Test Executives
- System Hardware
- Troubleshooting

---
RI ATE System Architecture
- Simple Direct Measurements
- Packaged or Wafer
[Speaker Notes: Both Cassini and the RI7100A systems are based on a single channel receiver design. The system uses simple test and measurement hardware to perform all measurements and relies on its object-oriented system software to control the measurement process, analyze the data, and provide a graphical user interface. The measurement process is simple and direct, signals to and from the device-under-test (DUT) are provided through several types of connection schemes: Independent On Wafer Coplanar Probes On Wafer Probe Cards or Membrane Probes A Device Contactor Assembly (socket) & DUT Interface board These contact devices can be connected to the universal RF Test Head of the RI 7100A through a hard dock test fixture or soft docked with individual cables.]

---
Cassini Architecture
- Infrastructure
  - System Rack (Large or Small)
  - System Power Supply
  - System Controller with RIFL III
  - RF Test Head (8 or 16), Fixture and DUT Interface
- Rack Modules
  - System Receiver
  - RF Sources
- Test Instrument Modules (TIMs)
  - Testset (banded from 6 GHz to 120 GHz)
  - Digital, DC and RF Instruments

---
Operating the System
- System power up sequence
- Using Fixture and DUT Board
- Docking with a Handler
- Using the System Controller (RIFL, eCS)
- Start and Stop RI System software
- Using Guru and the RI System Software

---
Starting and Stopping the System
- Check all power cables are connected
- Turn the MAIN POWER to ON
  - Switch on PCU & 48V Distribution Panel
- Turn on system controller and monitor(s)
- Wait 20 minutes
- Fixture should NOT be connected
- Emergency Off
[Speaker Notes: When first installing a RI Test System, make sure all the instruments in the rack are connected to the internal 48 Volt distribution block or the AC Power Conditioning Unit (PCU) or Uninterruptible Power Supply (UPS). Check that any external safety devices, such as circuit breakers or emergency off switches, are set to the "On" or "Operate" mode and that the main AC power supply cable(s) is plugged into the correct type of outlet. The system is configured at the factory for either 230 volt international standard AC power or US domestic 115 volt AC power. The system controller should always be turned on after or simultaneously with the main system rack, never before. The Microwave Sources and RFIC Test Head all contain heater ovens to maintain stable elevated temperatures within their critical components. The temperature will take a minimum of 20 minutes to stabilize within the units and we recommend a minimum of TWO HOURS warm up time before attempting a calibration or any precision measurements. Even if the power is shut off for only a very short amount of time, the ovens will corrupt any measurements while they go through a startup phase by increases the temperature to accelerate the warm-up process. ATTENTION! The 20 minute warm up time is required before running test plans, a minimum 2 hours before any calibration. The RF Test Fixture should not be connected when turning the system on for the first time. It could be possible to damage the Fixture because the System Controller has not yet reset the states of the Test Head. The fixture can be freely swapped while the system is idle or off, but a fixture should not be connected during startup.]

---
Power Supply Switches
Cassini 16
[Speaker Notes: To turn the system on, first move the switch labeled "MAIN POWER" to the "ON" position on the back of the system. From the front of the system, make sure the "Emergency Off" button is de-activated (popped out) and turn the front panel switch momentarily to 'START', then leave it set to the 'ON' position. After turning on the PCU or UPS, Cassini systems will automatically start the system controller and power on the monitor. For the RI7100A, the Monitor and System Controllers will need to be started up manually, usually by pressing the Power button on the front of the PC tower. To turn the system off, follow the RI System Software shut down procedure (described later), move the switch on the front panel to the 'OFF' position, then turn Off the PCU or UPS accessed from the back of the system and remove the Fixture from the Test Head. WARNING! Fixture should NOT be attached to the Test Head while powering on the system. The Power Conditioning Unit (PCU) is standard equipment for all current and future RI7100A and Cassini ATE systems. The power indicator light will illuminate when connected to the AC power mains. The 'Active' lights will glow green when power is being supplied to the system. Do not toggle the local/remote switch that is "locked down" to prevent tampering. If the "Emergency Off" button is pressed, all power to the system will be immediately disconnected, including the System Controller. Because the shut down procedure was not completed, some test data may be lost and those parts would have to be tested again.]

---
Power Supply Switches
Large Cassini
Small Cassini
[Speaker Notes: To turn the system on, first move the switch labeled "MAIN POWER" to the "ON" position on the back of the system. From the front of the system, make sure the "Emergency Off" button is de-activated (popped out) and turn the front panel switch momentarily to 'START', then leave it set to the 'ON' position. After turning on the PCU or UPS, Cassini systems will automatically start the system controller and power on the monitor. For the RI7100A, the Monitor and System Controllers will need to be started up manually, usually by pressing the Power button on the front of the PC tower. To turn the system off, follow the RI System Software shut down procedure (described later), move the switch on the front panel to the 'OFF' position, then turn Off the PCU or UPS accessed from the back of the system and remove the Fixture from the Test Head. WARNING! Fixture should NOT be attached to the Test Head while powering on the system. The Power Conditioning Unit (PCU) is standard equipment for all current and future RI7100A and Cassini ATE systems. The power indicator light will illuminate when connected to the AC power mains. The 'Active' lights will glow green when power is being supplied to the system. Do not toggle the local/remote switch that is "locked down" to prevent tampering. If the "Emergency Off" button is pressed, all power to the system will be immediately disconnected, including the System Controller. Because the shut down procedure was not completed, some test data may be lost and those parts would have to be tested again.]

---
Outline2
Chapter 1 Outline
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Guru Log-on, Applications
- Test Executives
- System Hardware
- Troubleshooting

---
Using the RF Fixture & DUT Board
[Speaker Notes: Low Cost, High Performance DUT Boards Full Integrated Calibration Complete Fixture Kit Parts In-stock Web Ordering Simple of Flex Docking Fastest Time to Manufacturing]

---
RF Fixture Layers

---
DUT Interface Board

---
Fixture Top Plate

---
Dual Site DUT Interface Board

---
DUT Board Frame

---
Fixture Module Expansion

---
Chapter 1 Outline
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Guru Log-on, Applications
- Test Executives
- System Hardware
- Troubleshooting

---
Socket in a Fixture

---
Docking with a Handler

---
Handler Docking Option

---
Docking Plate

---
Dual Site Fixture

---
Hander Interface Pod
- Plug and Play
- Configurable to work with any Handler
- Custom wiring by Customer

---
Chapter 1 Outline
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Guru Log-on, Applications
- Test Executives
- System Hardware
- Troubleshooting

---
Using System Controller and RIFL

---
System Controller Functions
- User Interface
- System Management
- Test Plan Generation & Execution
- Measurement Control and Signal Processing
- Data Analysis
[Speaker Notes: All of the user interface, system management, test plan generation, test plan execution, measurement control and data analysis is performed by the System Controller. The test system uses a high performance Intel processor-based server running OS/2, IBM’s multitasking operating system. The system software contains over 100 man years of high level executable software written in an object oriented programming language called SMALLTALK. Less than 1% of the software is written in assembly language for vector math.]

---
System Controller Contents
- x86 based Computer (embedded or tower)
- eCS Operating System
- RI System Software
- RIFL Bus (RI Fiber Link)
  - Three Generations (I,II,III)
[Speaker Notes: The RI Fiber Link (RIFL) found in early RI7100A systems used optical cables, but RIFL II uses shielded RJ45 cables (not typical "CAT5" network cables) and RIFL III uses a distictive round 10 pin connector. The RIFL ports are not ordered, so any device can be attached to any available port on the RIFL hub. To connect a RIFL cable, insert the RIFL II (shielded RJ45) or RIFL III (Round 10 pin) connector into an open RIFL port, until it clicks. To remove, press down on the tab along the top of the connector and pull. You can plug the cable into either the hub or the instrument first, the order does not matter. RIFL II to RIFL III cables are sometimes used to connect the different physical interfaces.]

---
RIFL II and III
- Cassini & RI7100A Gen 3 RI ATE System Communication and Control
- RIFL II and RIFL III connectors
- RI Interface Dongle plugs into System Computer's Parallel Port
- RI Instrument Control thru RIFL II Decoder Module
- External GPIB control through RIFL II to GPIB Interface Pod
- Plug and Play auto configuration of RIFL Nodes
- Scheduled Timing and Event Control with 1 µsec resolution
[Speaker Notes: For recent RI7100A and Cassini systems, the RI System interface and the RI Fiber Link 2 (RIFL II) provide the test system's internal low latency communication and instrumentation control link. The RI System interface state machine controller is an external dongle attached to the System Computer's parallel port. It provides the RIFL II interface connection and timing clock signal used throughout the system. The second generation RIFL "2" does not use fiber-optic cables any longer, but uses the efficient RIFL protocol with a proprietary bus that connects the System Interface to the RI instrumentation similar to a typical network hub configuration. RIFL II is not unshielded Category 5 (CAT5) cable typically used on computer network LANs. Each RI instrument contains a RIFL II interface connection and a RIFL II Decoder Module. RI instrumentation in the test system do not contain microprocessors for control. Control of all functions is provided by the System Computer through the RIFL II bus to the RIFL II Decoder in each RI instrument. A RIFL Decoder Module is an independent node on the RIFL II network bus. ATTENTION! RIFL II is not unshielded Category 5 (CAT5) cable typically used on computer network LANs. To control any GPIB instruments in the system, an RI System GPIB Pod contains a RIFL II Decoder Module, a RIFL II to GPIB Interface and GPIB connector for connecting GPIB cables between the RI System and any GPIB instruments. The RIFL II bus is a self addressing (plug and play) interface. The RIFL interface transfers serial data, and provides the system with 1 µsec scheduled timing and event control. For Cassini systems, the RJ45 interface is replaced by a distinctive round 10 pin connector that is electrically identical to the RIFL II connectors.]

---
History of RIFL
- 1st and 2nd Generation RI ATE System Communication and Control
- RI Fiber-optic Link (RIFL) Communication protocol
- RI Interface ISA PC Card Plugs into System Computer ISA Bus
- RI Instrument Control thru RIFL Decoder Modules
- External GPIB control through RIFL to GPIB Interface in the System Receiver
[Speaker Notes: The RI System Interface and the RI Fiber Link (RIFL) provide the test system’s optically isolated internal high bandwidth communication and instrumentation control link. The RI System Interface state machine controller uses standard ISA interface on the System Controller (computer). It provides the RIFL interface connection and timing clock signal. The RI Fiber Link uses the efficient RIFL protocol in a daisy chain, or token ring, LAN type proprietary bus that connects the System Interface to each of the RI instruments. Each RI instrument contains a RIFL interface connection and a RIFL Decoder Module. RI instrumentation in the test system does not contain microprocessors for control. Control of all functions is provided by the System Computer through the RIFL bus to the RIFL Decoder in each RI instrument. Each RIFL Decoder Module is an independent node on the RIFL network bus and must have a unique Node Number. To control a GPIB instruments in the system, an RI System GPIB connection is provided on the rear panel of the system receiver. The receiver contains a RIFL Decoder Module, a RIFL to GPIB Interface and GPIB connector for connecting GPIB cables between the Receiver and any GPIB instruments. The RIFL interface is unidirectional and transfers high speed serial data that provides the system with both precise scheduled timing (1 µsec resolution) and event control.]

---
Using OS/2 on the System Controller

---
System Software
- Graphical User Interfaces
- Production Package Part Test Executive
- Viewers and Data Saving
- Handler Control
- On Screen Help
[Speaker Notes: This is a microwave expert system, design by RF test engineer for RF test engineer. Graphical Programming Language ideal for production testing both package part and on-wafer testing flexible enough for engineering use. Extended tools to view and saving the test data handler control via handle pod on screen help]

---
Chapter 1 Outline
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Guru Log-on, Applications
- Test Executives
- System Hardware
- Troubleshooting

---
Start Up and Shut Down of RI System Software
To Log-on...
To Log-off...

---
Control with the Mouse and Keyboard
- System Control button
- Close button
- Max, Min buttons
- Window List

---
Mouse Basics
- Left Mouse Button
- Right Mouse Button

---
Mouse Shortcuts
- Copy RI Button ([CTRL] + Left)
- Select and Paste (Right)

---
OS/2 Utilities (Zip &Unzip)
- Long File Names & Extensions
- Zip/Unzip

---
Starting and Stopping RI System
- Click “Apps” or “Short Cuts” button
- System/Quit to exit
- Check that System is started
- Check that Fixture is docked

---
Chapter 1 Outline
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Guru Log-on, Applications
- Test Executives
- System Hardware
- Troubleshooting

---
On-Screen Help
Context sensitive help for each object
Searchable help system for browsing
Function key support for pointer location
User notes are definable for every panel
Hierarchical buttons with user definable notes

---
Testing Devices
  A.  Testplan defines "what" to test
1. Measurements performed
2. Values of instruments during test
3. Order of tests
4. Values of limits checked for pass/fail
  B.  Test Exec defines "how" to test it
1.  Handler control and interaction
2.  Operator user interface
3.  Definition of bins (soft and hard)
4. Logging of test data

---
Locating Test Plans

---
Test Executive Results

---
Wafer Probe Exec

---
Chapter 1 Outline
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Guru Log-on, Applications
- Test Executives
- System Hardware
- Troubleshooting

---
Cassini Architecture
- Infrastructure
  - System Rack (Large or Small)
  - System Power Supply
  - System Controller with RIFL II
  - RF Test Head, Fixture and DUT Interface
- Rack Modules
  - System Receiver
  - RF Sources
- Test Instrument Modules (TIMs)
  - Testset
  - DC and RF Instruments

---
Cassini Infrastructure
- Cassini Small or Large
- RIFL Hub
- System Controller

---
Test Head Configurations (TIMs)

---
Typical Testset TIM
- 4 ports
- External CW Synthesizer
- External DMSG
- External Receiver
- 20 GHz RF Testset
- Same Hardware
  as RI7100A RF

---
Typical Testset TIM

---
Cassini Receiver

---
Cassini Receiver

---
RF Sources
Functions:
- RF Signal Conditioning
- Combine RF lntermodulation Tones
- Test Head DC, Control Bus and RF Cable Routing
Contains:
- Six Plug-in RF Slots
- Attenuator/Amplifier/Combiner Module
- 0.01-6 GHz; Up to +10 dBm
- 0.1-4 GHz; Up to +22 dBm
- Auxiliary Source Amplitude Control Module
- GPIB Control Pod for external equipment
- RI High Speed Control Bus Hub (RIFL II)
- System Receiver

---
Source and Amp Attenuator

---
Source and Amp Attenuator

---
Device Power TIM
16 Bit DC & Pulsed Bias (Force & Sense)
  - Three Unipolar 3A VI's
  - Eight Bipolar 200mA VI's
  - Two, 3X8 Channel Matrix
16 Bit DUT Digital Control (Serial & Parallel)
  - Switchable 2mA Bipolar VI
12 Bit Voltage Measures Lines
  - Eight 50 K Samples/sec, Single or Differential
Analog Stimulus & Measurement (Including I & Q Tones)
  - 2 channel 10 or 40 MHz 12 Bit Arbs
  - 2 channel 90 MHz Digitizer
  - 40 MHz DDS Sine wave Source

---
Device Power / DUT Controller

---
RF Blind Mate Port Connections

---
Testset Block Diagram

---
Testset Internals

---
Receiver Diagram

---
Two Receivers in One

---
System Block Diagram

---
Chapter 1 Outline
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Guru Log-on, Applications
- Test Executives
- System Hardware
- Troubleshooting

---
System Troubleshooting

---
Handler Issues
- Handler cable wiring
- Handler Pod defines unique settings for handler
- Signals exchanged by handler and system
  - Start Test
  - End Test
  - Bin Part

---
Fixture Issues
- Auto-detection depends on serial number
- Fixture and DUT I/F have serial chip
- If both are new, Fixture must be "taught" before DUT IF
- Connector Hygiene

---
GPIB Instrument Issues
- GPIB cable length limit
- Don't extend GPIB cables to make a longer run
- Instrument made "inactive" if it fails at Startup
- GPIB instruments must have a unique address
- GPIB address must match stored address in Tester software object

---
LAB A
Report to Training Systems
Become comfortable operating RI System Software by running a Test Executive

---
Chapter Review
- System Power On and Off
- RF Fixture & DUT Board
- Docking with a Handler
- System Controller
- Running Test Plans
- Viewing Test Measurement Data
- System Hardware
- Troubleshooting

---
Preview Next Chapter Administration and Maintenance
- Overview of Guru, Applications
- Guru Log-on, User Privileges
- Guru Browser to Manage Guru Objects
- Guru Address Book, Update Guru Connections
- Software Updates and Patches
- System Networking and Data Logging
- Guru System Restore, Sync, and Log
- System Software Backups
- Guru Agents, Transferring Test Data
- System Software Errors & Recovery
- Monitor Error and Warning Messages
- Exchanging Cassini Modules
- Maintenance (Daily, Monthly, Annual PMC)

---
Questions?
Any Questions from this Chapter?