Module Installation RIK0127A
High Speed Digital
Output Header
M1 or M5 under
this corner
Mux Header
NOTE: This module will only
 work in the M1 or M5 position.

---
Header Designations
P1-P3 and P6 - P8 can be set for
clock, data, strobe, or static.
P9 - P16 = GND
 P1
P8
P2
P7
P3
P4
P5
P6
P4 = Read A
P5 = Read B
Mux In/Out
Any Pin
GND
Any of the input or output pins
can be routed to the mux pin.

---
Configuring the RIK0127A
High Speed Digital Module
1. Perform a system startup and confirm that
    the fixture is automatically activated.
2. Go to the main Test Environment Window
    and select “Test” and from the pull down
    menu “Equip...”.
3. When the Equipment Pool window
    opens, select “Control Modules”
    A “Module Browser” window will open.

---
Adding the Module
4. Scroll down the listing in the Module Browser
    and highlight “Y00065A1”.
5. Next go to the “Modules” selection on the
    main menu of this window and select
    “Add Module”.
6. A selection window will pop up. Select
    “Y2000H1A” and then “Select”.

---
Adding the Module
7. You will be prompted to provide
    the module location. Change the
    default location to “M5” if you are
    using the module in the M5 position.
    Leave it in the default state if you are
    using the M1 location on the carrier
    then select “OK” from the prompt.
NOTE: This module will only work in M1 or M5 locations!
8. Highlight the “M1 Y2000H1A” that will
    appear in the right hand column of the
    window. Several buttons will appear
    in the bottom half. Find the one named
    “Program”. Check to make sure it has
     Y18NA11A contained in its field.
Highlight

---
Saving The Module
Adding the Instrument
9. Highlight “Y00065A1” and “M1 Y2000H1A” in
    the Module Browser.Go to the Module Browser
    main menu and select “EE Save Modules”.
    Answer “Yes” to the prompt.
10. Highlight “Y00065A1” and “M1 Y2000H1A” in
      the Module Browser.
11. Go to the Module Browser main menu and
      select “Node” and then “Add Inst”.
12. Select “RI74315C” and then “select”.

---
Adding the Instrument
12. When prompted for the “Instrument Name”
      choose “OK”. This will allow the default
      “FixtureDigital1” to be entered.
13. You will then be prompted to enter the
      “Module Location Integer”. Enter 1 for M1
       or 5 for M5 and select “OK”
NOTE: This module will only work in M1
        or M5 locations!
14. Highlight “Y00065A1” and “M1 Y2000H1A” in
      the Module Browser.
15. Go to the Module Browser main menu and
      select “Node” and then “EE Save Node Info”
      from the pull down menu. Respond “Yes” to the
      prompt.

---
Checking The Changes
15. Close the Module Browser and the equipment pool windows
      and then perform a system
      startup.
16. Deactivate the fixture and
      perform a system startup.
      Return to the Module Browser
      window and observe to see if
      the full configuration has been
      saved.
Module and
    Location
Carrier and Instrument
Number
Program Name
Location/Name
 are equal

---
Module Buttons
Device DB
Set On/Off
 Voltages
Set to midpoint
 between Von and Voff
Db1-3 ans Db6-8
can be set to these functions.

---
Module Buttons
Device Serial Patterns
Clock cycles before
 read start
For local writes

---
Module Buttons
Scan
Results of scan on
 ReadA (Db4)
Results of scan on
 ReadB (Db5)
Clock pulses to skip before read

---
Displaying Scan Results
Compare Enable
Comparator Out
Fail Bit
Expected Out
Display Scan Out Example
Set to 500n sec in
this example

---
Preparing to Display Scan Results
Measurement
Scan file compared to
Clock pulses to skip before read.
Used to compensate for DUT signal
delay.
Note:
Before the Scan Results
can be displayed “Cmp Scan”
must be performed

---
Measure and Display Fail Data
Return the results based on ReadA (Db4).
Assumes Cmp Scan A was done.
Traces are defined as:
    Bottom = Fail Bit
    Second from Bottom = Read Expected
    Third from Bottom = Comparator Output
    Fourth from Bottom = Compare Enable.
Make sure to use a line type of “binary”

---
Loading a Scan File1
Once the carrier board is configured a scan file can be loaded into the
modules memory. Here are the steps required to do that.
Go to the Tester Configuration window and
highlight the fixture digital.
Choose the “Load Patterns” from the button
list.
Choose the .plist file that
   represents the scan file(s) that
   you want to load.

---
Loading a Scan File2
Once you have selected the pattern to load the tester will ask if you
   want to erase existing patterns.
You will then need to wait for the eraser and loading of the new patterns.

---
Plist Example
Company or product name (annotation)
Line
Remove “,” and change to upper case (defines default behavior
1
2
3
4
5
6
7
8
9
10
title=XXXXXX scan loader
parse=(vector riRemoveChar:$,)asUpperCase
skip=(vector size > 3)
signalTypes,TDS=scanStrobe,TDI=scanDrive,TDO=scanExpect
signals=TDS;TDI;TDO
allowed=10X;11X;01X;00X;000;001;010;011;100;101
translate=10X=SdC;11X=SDC;01X=sDC;00X=sdC;000=sdrC;001=sdRC;
010=sDrC;011=sDRC;100=SdrC;101=SDrC
comment=pattern name,file name, bytes per vector,version
scanFile,name=shiftDsp,file=shiftdsp3.csv,vectorSize=7,version=1.0
 Scan File CSV
-- /TDS /TDI /TDO
1,0,x
1,0,x
1,0,x
1,0,x
1,0,x
Ignore any line more than 3 long
Signal types allowed.
Can also have TCK = scanClock and TRST=scanAux
Signals present and column order
0,0,x
0,0,0
0,0,0
0,1,1
0,0,0
etc..
...
..
What patterns are allowed and how they are
translated.
Annotation
3 vectors + 2 ”,” + CR + LF