Intro
Roos Instruments, Inc.
CASSINI
RF ATE System
Operator Training

---
Outline1
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
RI ATE System Architecture
Simple Direct Measurements
Packaged or Wafer
[Speaker Notes: Both Cassini and the RI7100A systems are based on a single channel receiver design. The system uses simple test and measurement hardware to perform all measurements and relies on its object-oriented system software to control the measurement process, analyze the data, and provide a graphical user interface. The measurement process is simple and direct, signals to and from the device-under-test (DUT) are provided through several types of connection schemes: Independent On Wafer Coplanar Probes On Wafer Probe Cards or Membrane Probes A Device Contactor Assembly (socket) & DUT Interface board These contact devices can be connected to the universal RF Test Head of the RI 7100A through a hard dock test fixture or soft docked with individual cables.]

---
Cassini Architecture
Infrastructure
  - System Rack (Large or Small)
  - System Power Supply
  - System Controller with RIFL II
  - RF Test Head, Fixture and DUT Interface
Rack Modules
  - System Receiver
  - RF Sources
Test Instrument Modules (TIMs)
  - Testset
  - DC and RF Instruments

---
Starting and Stopping the System
Check all power cables are connected
Turn the MAIN POWER to ON
  - Switch on PCU or UPS
Turn on system controller and monitor(s)
Wait 20 minutes
Fixture should NOT be connected
Emergency Off
[Speaker Notes: When first installing a RI Test System, make sure all the instruments in the rack are connected to the internal 48 Volt distribution block or the AC Power Conditioning Unit (PCU) or Uninterruptible Power Supply (UPS). Check that any external safety devices, such as circuit breakers or emergency off switches, are set to the "On" or "Operate" mode and that the main AC power supply cable(s) is plugged into the correct type of outlet. The system is configured at the factory for either 230 volt international standard AC power or US domestic 115 volt AC power. The system controller should always be turned on after or simultaneously with the main system rack, never before. The Microwave Sources and RFIC Test Head all contain heater ovens to maintain stable elevated temperatures within their critical components. The temperature will take a minimum of 20 minutes to stabilize within the units and we recommend a minimum of TWO HOURS warm up time before attempting a calibration or any precision measurements. Even if the power is shut off for only a very short amount of time, the ovens will corrupt any measurements while they go through a startup phase by increases the temperature to accelerate the warm-up process. ATTENTION! The 20 minute warm up time is required before running test plans, a minimum 2 hours before any calibration. The RF Test Fixture should not be connected when turning the system on for the first time. It could be possible to damage the Fixture because the System Controller has not yet reset the states of the Test Head. The fixture can be freely swapped while the system is idle or off, but a fixture should not be connected during startup.]

---
Pictures of Power Switches
[Speaker Notes: To turn the system on, first move the switch labeled "MAIN POWER" to the "ON" position on the back of the system. From the front of the system, make sure the "Emergency Off" button is de-activated (popped out) and turn the front panel switch momentarily to 'START', then leave it set to the 'ON' position. After turning on the PCU or UPS, Cassini systems will automatically start the system controller and power on the monitor. For the RI7100A, the Monitor and System Controllers will need to be started up manually, usually by pressing the Power button on the front of the PC tower. To turn the system off, follow the RI System Software shut down procedure (described later), move the switch on the front panel to the 'OFF' position, then turn Off the PCU or UPS accessed from the back of the system and remove the Fixture from the Test Head. WARNING! Fixture should NOT be attached to the Test Head while powering on the system. The Power Conditioning Unit (PCU) is standard equipment for all current and future RI7100A and Cassini ATE systems. The power indicator light will illuminate when connected to the AC power mains. The 'Active' lights will glow green when power is being supplied to the system. Do not toggle the local/remote switch that is "locked down" to prevent tampering. If the "Emergency Off" button is pressed, all power to the system will be immediately disconnected, including the System Controller. Because the shut down procedure was not completed, some test data may be lost and those parts would have to be tested again.]

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Using the RF Fixture & DUT Board
[Speaker Notes: Low Cost, High Performance DUT Boards Full Integrated Calibration Complete Fixture Kit Parts In-stock Web Ordering Simple of Flex Docking Fastest Time to Manufacturing]

---
RF Test Fixture Layers

---
DUT Interface Board

---
Fixture Top Plate

---
Dual Site DUT Interface Board

---
DUT Board Frame

---
Fixture Module Expansion

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Socket in a Fixture

---
Docking with a Handler

---
Handler Docking Option

---
Docking Plate

---
Dual Site Fixture

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Using System Controller and RIFL

---
System Controller Functions
User Interface
System Management
Test Plan Generation & Execution
Measurement Control and Signal Processing
Data Analysis
[Speaker Notes: All of the user interface, system management, test plan generation, test plan execution, measurement control and data analysis is performed by the System Controller. The test system uses a high performance Intel processor-based server running OS/2, IBM’s multitasking operating system. The system software contains over 100 man years of high level executable software written in an object oriented programming language called SMALLTALK. Less than 1% of the software is written in assembly language for vector math.]

---
System Controller Contents
x86 based Computer (embedded or tower)
OS/2 Operating System
RI System Software
RIFL Bus (RI Fiber Link)
  - Three Generations (I,II,III)
[Speaker Notes: The RI Fiber Link (RIFL) found in early RI7100A systems used optical cables, but RIFL II uses shielded RJ45 cables (not typical "CAT5" network cables) and RIFL III uses a distictive round 10 pin connector. The RIFL ports are not ordered, so any device can be attached to any available port on the RIFL hub. To connect a RIFL cable, insert the RIFL II (shielded RJ45) or RIFL III (Round 10 pin) connector into an open RIFL port, until it clicks. To remove, press down on the tab along the top of the connector and pull. You can plug the cable into either the hub or the instrument first, the order does not matter. RIFL II to RIFL III cables are sometimes used to connect the different physical interfaces.]

---
RIFL II and III
Cassini & RI7100A Gen 3 RI ATE System Communication and Control
RIFL II and RIFL III connectors
RI Interface Dongle plugs into System Computer's Parallel Port
RI Instrument Control thru RIFL II Decoder Module
External GPIB control through RIFL II to GPIB Interface Pod
Plug and Play auto configuration of RIFL Nodes
Scheduled Timing and Event Control with 1 µsec resolution
[Speaker Notes: For recent RI7100A and Cassini systems, the RI System interface and the RI Fiber Link 2 (RIFL II) provide the test system's internal low latency communication and instrumentation control link. The RI System interface state machine controller is an external dongle attached to the System Computer's parallel port. It provides the RIFL II interface connection and timing clock signal used throughout the system. The second generation RIFL "2" does not use fiber-optic cables any longer, but uses the efficient RIFL protocol with a proprietary bus that connects the System Interface to the RI instrumentation similar to a typical network hub configuration. RIFL II is not unshielded Category 5 (CAT5) cable typically used on computer network LANs. Each RI instrument contains a RIFL II interface connection and a RIFL II Decoder Module. RI instrumentation in the test system do not contain microprocessors for control. Control of all functions is provided by the System Computer through the RIFL II bus to the RIFL II Decoder in each RI instrument. A RIFL Decoder Module is an independent node on the RIFL II network bus. ATTENTION! RIFL II is not unshielded Category 5 (CAT5) cable typically used on computer network LANs. To control any GPIB instruments in the system, an RI System GPIB Pod contains a RIFL II Decoder Module, a RIFL II to GPIB Interface and GPIB connector for connecting GPIB cables between the RI System and any GPIB instruments. The RIFL II bus is a self addressing (plug and play) interface. The RIFL interface transfers serial data, and provides the system with 1 µsec scheduled timing and event control. For Cassini systems, the RJ45 interface is replaced by a distinctive round 10 pin connector that is electrically identical to the RIFL II connectors.]

---
History of RIFL
1st and 2nd Generation RI ATE System Communication and Control
RI Fiber-optic Link (RIFL) Communication protocol
RI Interface ISA PC Card Plugs into System Computer ISA Bus
RI Instrument Control thru RIFL Decoder Modules
External GPIB control through RIFL to GPIB Interface in the System Receiver
[Speaker Notes: The RI System Interface and the RI Fiber Link (RIFL) provide the test system’s optically isolated internal high bandwidth communication and instrumentation control link. The RI System Interface state machine controller uses standard ISA interface on the System Controller (computer). It provides the RIFL interface connection and timing clock signal. The RI Fiber Link uses the efficient RIFL protocol in a daisy chain, or token ring, LAN type proprietary bus that connects the System Interface to each of the RI instruments. Each RI instrument contains a RIFL interface connection and a RIFL Decoder Module. RI instrumentation in the test system does not contain microprocessors for control. Control of all functions is provided by the System Computer through the RIFL bus to the RIFL Decoder in each RI instrument. Each RIFL Decoder Module is an independent node on the RIFL network bus and must have a unique Node Number. To control a GPIB instruments in the system, an RI System GPIB connection is provided on the rear panel of the system receiver. The receiver contains a RIFL Decoder Module, a RIFL to GPIB Interface and GPIB connector for connecting GPIB cables between the Receiver and any GPIB instruments. The RIFL interface is unidirectional and transfers high speed serial data that provides the system with both precise scheduled timing (1 µsec resolution) and event control.]

---
Using OS/2 on the System Controller

---
System Software
Graphical User Interfaces
Production Package Part Test Executive
Production Wafer Test Executive
Viewers and Data Saving
Handler Control
On Screen Help
[Speaker Notes: This is a microwave expert system, design by RF test engineer for RF test engineer. Graphical Programming Language ideal for production testing both package part and on-wafer testing flexible enough for engineering use. Extended tools to view and saving the test data handler control via handle pod on screen help]

---
Start Up and Shut Down of RI System Software

---
Starting and Stopping RI System
Double-click icon to start
System/Quit to exit
Check that System is started
Check that Fixture is docked

---
System Logon Screen
RI Message Window
Software Version
System / Logon

---
System Users Window

---
Main Window

---
Shut Down System Software

---
Control with the Mouse and Keyboard
System Control button
Close button
Max, Min buttons
Window List

---
Mouse Basics
Left Mouse Button
Right Mouse Button

---
Mouse Shortcuts
Copy RI Button ([CTRL] + Left)
Select and Paste (Right)

---
OS/2 Utilities (Zip &Unzip)
Long File Names & Extensions
Zip/Unzip

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Running Test Plans

---
Wafer Probe Exec

---
On Screen Help
Context sensitive help for each object
Searchable help system for browsing
Function key support for pointer location
User notes are definable for every panel
Hierarchical buttons with user definable notes

---
Activating Tester and Handler

---
Active Hardware vs Simulation

---
Test Executive Container Window

---
Package Test Executives

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Viewing Test Measurement Data
Colored Last Test Window
Worksheet Based Detail and Stats Views
Graphical Test Plan Viewer
Bar Chart Viewers
Multi-Histogram Viewer

---
Worksheet Views

---
Bar Chart Viewers

---
Histogram Viewer

---
Test Exec Reports

---
Test Exec Failure Messages
Normal messages
Error messages
Software error display
Service Request

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Cassini Test Head Configurations

---
Cassini Instruments – Testset

---
Cassini Receiver

---
Source and Amp Attenuator

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
System Troubleshooting

---
Handler Issues
Handler cable wiring
Handler Pod defines unique settings for handler
Signals exchanged by handler and system
  - Start Test
  - End Test
  - Bin Part

---
Fixture Issues
Auto-detection depends on serial number
Fixture and DUT I/F have serial chip
If both are new, Fixture must be "taught" before DUT IF
Connector Hygiene

---
GPIB Instrument Issues
GPIB cable length limit
Don't extend GPIB cables to make a longer run
Instrument made "inactive" if it fails at Startup
GPIB instruments must have a unique address
GPIB address must match stored address in Tester software object

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Outline6
Chapter Outline (Day 2)
Creating and Editing User Log-ons
Software Revisions and Patches
System Networking and Data Logging
System Software Backups
System Software Errors & Recovery
Message Pop-up window
Editing Test Executives
Handler Bin Definitions
Test Exec Summary Reports
Data Saving & Releasing Test Execs
Exchange Cassini Modules
Preventive Maintenance (Daily,Weekly,Monthly)
Bi-Annual Preventive Maintenance and Calibration

---
System Software Errors & Recovery
Execution Halt
System is not responding
System is not responding to CTRL + Break
OS/2 is not responding to CTRL+ESC
Requesting Support

---
Errors that Halt Progress
System software Error display
Click “Save” to build vRtError.log file
located in D:\RiApps\Testexec
Perform a System “Reset”
Create a Service Request (File | Service Request)

---
Software Recovery – First Attempt
System is not responding
Press and hold CTRL key, press Break key
RI System will halt and display an Error dialog box
Click Abort to close the Error dialog box
RI Message Window, Select System, and Reset
Attempt to run your application again, if the error persists, click “Save”, then File, Service Request.
Send “.zip” file to support@roos.com
Select System, then Quit, Yes to close RI System Software (and restart)

---
Software Recovery – Second Try
System is not responding to CTRL + Break
Press & Hold CTRL Key and press ESC key
OS/2 will display the Window List (it may take several seconds)
Right click on “RI RFIC Test Environment” and select Close
Select Yest to close the RI System Software
Restart RI System Software

---
Software Recovery – Last Choice
OS/2 System is not responding to CTRL + ESC
Press and hold CTRL + ALT + DELETE Key
This will REBOOT the system controller
Restart the system controller software
note: The system always saves to disk each part's test data after each test run has been completed.

---
Documenting Software Issues
Description of issue found
Procedure to duplicate the issue
Text copies of any error messages and vRtError.log
Use the Service Request feature to send a .zip file containing all needed files to duplicate the error

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Message Pop-up Window
The Message window appears as “Log File”

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Binning Control
Set up “Special Fails” (also called Categories)
If using handler with “retest” set up retest criteria
Choose handler bins for pass, fail, and special fails.

---
Set up Special Fails
Select “Special Fails” from “Test” menu

---
Special Fail Criteria
Map the Special Fail Category to failed tests
Rename the category by clicking the Rename button

---
Set up Retest Criteria
Select “Retest Criteria” from the “Test” menu
Choose one or more criteria
Select Retry Count

---
Set up Pause Criteria
Select “Pause Criteria” from “Test” menu
Set the condition to pause the tester

---
Set up Handler Bins
Select “Edit Handler Bins” from the “Options” menu

---
Set up Handler Bins

---
Set Up Handler Bins
Choose “Hard” and “Soft” bins for each category
Hard bin is the sort value sent to the handler
Soft bin is stored in data logs such as STDF
If using Retest Criteria, you should use a handler capable of “re-plunge”
Set the “retest bin” signal sent to the handler to cause the re-plunge of the part

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Summary Reports

---
Reset Lot Statistics

---
Set “Abort Test on Fail”
Ends Primary test plan as soon as a test fails
Only aborts a test section boundary
Internally “skips” across each section
Typically do “DC Continuity” test in the first section
“Special fails” can be affected when later tests are skipped

---
Choose Views to Open
Set one or more views to open automaticaly
When Test Executive starts, windows are opened
Window position and size are not saved

---
Set Preference for “Avg Yield” or “Avg Throughput”
Average Yield is 35 point moving average
Average Throughout uses 50 point moving average of test time

---
Release Test Executive
“Release” locks the Test Exec from further changes (makes it read-only)
Can be made writable only by copying
Plan on version control when naming Test Execs

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual)

---
Cassini TIM Removal
Cassini Test Instrument Module (TIM)
Can be removed with power on
Remove by pulling release handles

---
Cassini TIM Insertion
Can be inserted while the system is running
Align module in Test Head, Push up until you hear a click, then push in top handle to power on the module

---
Outline
System Power On and Off
RF Fixture & DUT Board
Docking with a Handler
System Controller
Running Test Plans
Viewing Test Measurement Data
System Hardware
Troubleshooting
Software Errors & Recovery
Message Popup Window
Handler Bin Definitions
Test Exec Summary Reports
Exchange Cassini Modules
Maintenance (Daily, Weekly, Monthly, Bi-Annual PMC)

---
Preventive Maintenance Schedule
Daily
Weekly
Monthly
Bi-Annually

---
Daily Preventive Maintenance
Visually inspect the Test Socket, Load Board, and connectors
Visually inspect all of the RF Connectors around the Test Fixture
Visually inspect the Docking Hardware
Run known good and bad parts

---
Weekly Preventive Maintenance
Perform Daily PMC
Visually inspect all the RF connectors around the Test Head Pogo Ring
Visually inspect RIFL cable to Handler Pod and Handler interface

---
Monthly Preventive Maintenance
Perform Weekly PM
Perform the “System Verify Procedure”

---
Bi-annual Preventive Maintenance and Calibration
Carefully inspect all the RF connectors around the test head Pogo Ring
Inspect all the cable connections and check all SMA cable connections for tightness
Clean the air filters on the back of all the RF Sources
Perform the Test Head Fixture Cam Assembly PM
Perform the “Procedure for Complete Cassini System Cal”

---
LAB A
Report to Training Laptops
Become comfortable operating RI System Software by running a Test Executive

---
Training Review
System Power On and Off (page 4)
RF Fixture & DUT Board (7)
Docking with a Handler (14)
System Controller (18)
Running Test Plans (35)
Viewing Test Measurement Data (47)
System Hardware (53)
Troubleshooting (62)
Software Errors & Recovery (66)
Message Popup Window (71)
Handler Bin Definitions (72)
Test Exec Summary Reports (78)
Exchange Cassini Modules (80)
Maintenance (Daily, Weekly, Monthly, Bi-Annual PMC) (81)

---
Questions?
Any Questions from this Chapter?